## Sheet4
NX2 | L-C-R | Exercise to determine Ipeak (for calibration of current coil), short- circuit inductance and stray resistance
Given: Co=28uF and Vo=11kV and the discharge current waveform; with accurate time base
Short circuit data |  | Method: From waveform, measure T; measure reversal ratio f=In+1/In
time us | current kA | Then: |  | Lo=T2/(4p2Co)
Io=pCoVo(1+f)/T
ro=-(2/p)Ln(f)(Lo/Co)0.5
From the waveform below, measure 3T and peak amplitudes and fill in cells with red ? below
|  | Co | Vo | Co*Vo
|  | 2.8e-05 | 11000 | 0.308
|  | Measure T, from |  | 3*T | T |  | Calculate: | Lo
|  | current waveform: |  | ? | #VALUE! |  | Lo | #VALUE! | [Calculated from Lo=T2/(4*p*p*Co}, ]
|  | Measure reversal ratio |  | I1 | I2 | I3 | I4 | I5 | I6 |  | Calculate | Ln(f) | Zo | ro
|  | f, from waveform | [8 empty] | ro | #VALUE! | #VALUE! | #VALUE! | bar: I1 to I6
|  | f1 | f2 | f3 | f4 | f5 | f
|  | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
|  |  |  | Calculate | Ipeak
[5 empty] | #VALUE! | [Calculated from low damping approximation: p*Co*Vo*(1+f)/T ]
From the above, fill in blue ? Below for comparison:
|  | Comparison: | [7 empty] | Lo | ro |  | Io | bar: I5 to I6
|  | Actual Values |  | I1: Co,Vo, Lo and ro are input values, Io resultant value; I5: 28; I6: 11 | 15 | 2.2 |  | 442
|  | Deduced from waveform, low damping approximation: | [7 empty] | ? | ? |  | ?
|  | Lo by measuring T: |  | I1: using formula Lo=T2/(4p2Co)
|  | f by measuring f1 to f5: |  | I1: using formula f=(1/5)(f1+f2+f3+f4+f5); I5: where f1=I2/I1; f2=I3/I2, f3=I4/I3, f4=I5/I4, f5=I6/I5
|  | Io by measuring T & f: |  | I1: using formula Io=pCoVo(1+f)/T
|  | ro by measurinf f: |  | I1: using formula ro=-(2/p)Ln(f)(Lo/Co)0.5
|  | Notes: | For integrity of baseline data; it is important that a plasma focus should have a reliable value of Co.
For a plasma focus capacitor bank, Co can be measured by a RC leak method
using a resistor of known value R.
Charge up Co (not necessarily to the usual operational voltage, a lower voltage will do)
Discharge Co through R, whilst monitoring the voltage across the capacitor.
Measure the time the voltage takes to drop from any value V1 to (1/e)V1 where
e=2.71828 is the natural exponential. This measured time is equated to RCo to
obtain the value of Co
Such a measurement has to be carefully designed with all the usual safety precautions.
The circuit and connections have to be carefully worked out; the resistor power and voltage ratings
considered before hand. We have found an electrostatic voltmeter very useful for such work,
enabling using a stop-watch for timing the RC.
You could of course use your preferred method.
Once the value of Co is fixed; then Lo and ro may be obtained from a discharge waveform.
Ensure that the oscilloscope has a calibrated time base. The value of Io for the discharge
can also be obtained from a measurement of T and f, as in the example above,
so that the coil monitoring that waveform can then be calibrated in situ.
If the capacitor bank and circuit is too heavily damped, the above method will not be accurate.
Then the fitting method has to be used. The computed current waveform is compared to the measured;
Lo and ro are adjusted until the waveform achieve a good fit.

## Sheet3
NX2 NEP35V110L50 |  | Lo nH | Co uF | ro | b cm | a cm | zo cm | Vo | Po torr Neon |  | Source: NTU/NIE Plasma Lab | [6 empty] | Neon gas(NEP35V110L50) |  | Po Torr | Ipeak | Ipinch | Peak va | S | Peak vs | Peak vp | amin | zmax | pnch dur | Vmax | T pinch | ni pinch | z | EINP % | Ysxr | ALT
|  | 15 | 28 | 2.2 mOhm | 4.1 | 1.9 | 5 | 11 | 2.6 |  | fm=0.1 | fc=0.7 | fmr=0.12 | fcr=0.68 | [6 empty] | kA | kA | cm/us |  | cm/us | cm/us |  |  | ns | kV | (10^6)K | (1023)/m3
Measured 2.6 Torr |  | Computed 2.6 Torr |  | Computed 1000000 Torr |  | Computed 4 Torr |  |  |  | Computed 0.5 Torr |  | Computed 2.9 Torr |  | Computed 1.5 Torr |  | Computed 5 Torr |  | L0=15 (nominal) |  | 1000000 | 442 | [9 empty] | average
time us | current kA | time us | current kA | time us | current kA | time us | current kA |  |  | time Us | current kA | time us | current kA | time us | current kA | time us | current kA | C0=28 Micro Farad | can't force | 10 | unable to run
-0.57 | 0 | 0.0012961 | 0.95051 | [8 empty] | 0.0012961 | 0.95051 | 0.0012961 | 0.95051 | 0.0012961 | 0.95051 | b=4.1cm | can force | 5 | 383 | 84 | 4.3 | 90 | 11 | 8.6 | 0.25 | 2.7 | 78.6 | 6.9 | 0.4 | 4.2 | 5.5 | 18.5 | 0.46 | 0.61 | pinch expanding
-0.565 | -0.06925 | 0.0025923 | 1.9008 | [8 empty] | 0.0025923 | 1.9008 | 0.0025923 | 1.9008 | 0.0025923 | 1.9008 | a=1.9cm | can force | 4.5 | [16 empty] | pinch expanding
-0.56 | 0.05351 | 0.0038884 | 2.851 | [8 empty] | 0.0038884 | 2.851 | 0.0038884 | 2.851 | 0.0038884 | 2.851 | Length of Anode=5.0cm | can run | 4 | [16 empty] | pinch not compressive
-0.555 | -0.11174 | 0.0051846 | 3.8009 | [8 empty] | 0.0051846 | 3.8009 | 0.0051846 | 3.8009 | 0.0051846 | 3.8009 | Circuit Resistance=2.2 mOhms |  | 3.5 | [16 empty] | pinch compressive
-0.55 | -0.08498 | 0.0064807 | 4.7506 | [8 empty] | 0.0064807 | 4.7506 | 0.0064807 | 4.7506 | 0.0064807 | 4.7506 | Mass Factor=0.1 |  | 3.2 | [16 empty] | pinch is compressive}
-0.545 | -0.05823 | 0.0077769 | 5.7001 | [8 empty] | 0.0077769 | 5.7001 | 0.0077769 | 5.7001 | 0.0077769 | 5.7001 | Current Factor=0.7 |  | 3 | [16 empty] | for all lower pressures}
-0.54 | 0.16052 | 0.009073 | 6.6494 | [8 empty] | 0.009073 | 6.6494 | 0.009073 | 6.6493 | 0.009073 | 6.6494 | Specific heat Ratio=1.4 |  | 2.9 | 369.4 | 145 | 5.8 | 114 | 20.5 | 14.5 | 0.16 | 2.79 | 30.2 | 20.3 | 1.63 | 5.7 | 8 | 27.5 | 22.6 | 0.8
-0.535 | 0.09128 | 0.010369 | 7.5984 | [8 empty] | 0.010369 | 7.5985 | 0.010369 | 7.5984 | 0.010369 | 7.5985 | Capacitor Voltage =11.0kV |  | 2.8
-0.53 | 0.21403 | 0.011665 | 8.5472 | [8 empty] | 0.011665 | 8.5473 | 0.011665 | 8.5471 | 0.011665 | 8.5474 | Filling Pressure in torr=2.6 |  | 2.7
-0.525 | 0.24079 | 0.012961 | 9.4958 | [8 empty] | 0.012961 | 9.4958 | 0.012961 | 9.4957 | 0.012961 | 9.496 | Mass factor in Radial Phase=0.12 |  | 2.6 | 366 | 152 | 6.1 | 120 | 22.5 | 15.3 | 0.18 | 2.8 | 27.9 | 21.5 | 2 | 4 | 8 | 27.8 | 15.8 | 0.85
-0.52 | 0.26754 | 0.014258 | 10.4441 | [8 empty] | 0.014258 | 10.4442 | 0.014258 | 10.4439 | 0.014258 | 10.4443 | Current Factor, radial=0.68 |  | 2.4
-0.515 | 0.1983 | 0.015554 | 11.3922 | [8 empty] | 0.015554 | 11.3922 | 0.015554 | 11.3919 | 0.015554 | 11.3924 | Gas,Atomic_no.,Atomic_mass=Neon,10,20 |  | 2
-0.51 | 0.32105 | 0.01685 | 12.34 | [8 empty] | 0.01685 | 12.34 | 0.01685 | 12.3396 | 0.01685 | 12.3403 | SXR yield=20J |  | 1.5 | 350 | 168 | 7.6 | 151 | 27.6 | 18.9 | 0.25 | 2.75 | 21.7 | 26.7 | 4.2 | 1.2 | 8.7 | 27.3 | 1.5 | 1.12
-0.505 | 0.3478 | 0.018146 | 13.2875 | [8 empty] | 0.018146 | 13.2876 | 0.018146 | 13.287 | 0.018146 | 13.2879 |  |  | 1
-0.5 | 0.37456 | 0.019442 | 14.2347 | [8 empty] | 0.019442 | 14.2348 | 0.019442 | 14.2341 | 0.019442 | 14.2352 |  |  | 0.5
-0.495 | 0.30531 | 0.020738 | 15.1817 | [8 empty] | 0.020738 | 15.1818 | 0.020738 | 15.181 | 0.020738 | 15.1823
-0.49 | 0.33207 | 0.022035 | 16.1283 | [8 empty] | 0.022035 | 16.1285 | 0.022035 | 16.1275 | 0.022035 | 16.1291 |  |  | Po Torr | Ipeak | Ipinch | Peak va | S | Peak vs | Peak vp | amin | zmax | pnch dur | Vmax | T pinch | ni pinch | z | EINP % | Ysxr J | ALT
-0.485 | 0.35882 | 0.023331 | 17.0747 | [8 empty] | 0.023331 | 17.0748 | 0.023331 | 17.0736 | 0.023331 | 17.0756 |  |  |  | [ Normalised to the value at optimum 2.9 Torr ]
-0.48 | 0.28957 | 0.024627 | 18.0207 | [8 empty] | 0.024627 | 18.0209 | 0.024627 | 18.0195 | 0.024627 | 18.0217 |  |  | 1000000
-0.475 | 0.41233 | 0.025923 | 18.9664 | [8 empty] | 0.025923 | 18.9666 | 0.025923 | 18.965 | 0.025923 | 18.9676 |  |  | 10
-0.47 | 0.63108 | 0.027219 | 19.9118 | [8 empty] | 0.027219 | 19.9121 | 0.027219 | 19.9102 | 0.027219 | 19.9132 |  |  | 5 | 1.0368 | 0.57931 | 0.74138 | 0.78947 | 0.53659 | 0.5931 | 1.5625 | 0.96774 | 2.6026 | 0.3399 | 0.2454 | 0.73684 | 0.6875 | 0.67273 | 0.020354 | 0.7625
-0.465 | 0.36984 | 0.028515 | 20.8569 | [8 empty] | 0.028515 | 20.8572 | 0.028515 | 20.855 | 0.028515 | 20.8585 |  |  | 4.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
-0.46 | 0.39659 | 0.029811 | 21.8016 | [8 empty] | 0.029811 | 21.8019 | 0.029811 | 21.7994 | 0.029811 | 21.8035 |  |  | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
-0.455 | 0.32734 | 0.031108 | 22.746 | [8 empty] | 0.031108 | 22.7464 | 0.031108 | 22.7435 | 0.031108 | 22.7481 |  |  | 3.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
-0.45 | 0.4501 | 0.032404 | 23.69 | [8 empty] | 0.032404 | 23.6904 | 0.032404 | 23.6872 | 0.032404 | 23.6924 |  |  | 3.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
-0.445 | 0.38085 | 0.0337 | 24.6336 | [8 empty] | 0.0337 | 24.6342 | 0.0337 | 24.6305 | 0.0337 | 24.6364 |  |  | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
-0.44 | 0.40761 | 0.034996 | 25.5769 | [8 empty] | 0.034996 | 25.5775 | 0.034996 | 25.5734 | 0.034996 | 25.58 |  |  | 2.9 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
-0.435 | 0.43436 | 0.036292 | 26.5198 | [8 empty] | 0.036292 | 26.5205 | 0.036292 | 26.5159 | 0.036292 | 26.5233 |  |  | 2.8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
-0.43 | 0.26911 | 0.037588 | 27.4624 | [8 empty] | 0.037588 | 27.4631 | 0.037588 | 27.458 | 0.037588 | 27.4662 |  |  | 2.7 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
-0.425 | 0.39187 | 0.038884 | 28.4045 | [8 empty] | 0.038884 | 28.4053 | 0.038884 | 28.3997 | 0.038884 | 28.4088 |  |  | 2.6 | 0.9908 | 1.0483 | 1.0517 | 1.0526 | 1.0976 | 1.0552 | 1.125 | 1.0036 | 0.92384 | 1.0591 | 1.227 | 0.70175 | 1 | 1.0109 | 0.69912 | 1.0625
-0.42 | 0.41862 | 0.040181 | 29.3463 | [8 empty] | 0.040181 | 29.3472 | 0.040181 | 29.341 | 0.040181 | 29.351 |  |  | 2.4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
-0.415 | 0.34938 | 0.041477 | 30.2876 | [8 empty] | 0.041477 | 30.2886 | 0.041477 | 30.2818 | 0.041477 | 30.2928 |  |  | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
-0.41 | 0.47213 | 0.042773 | 31.2286 | [8 empty] | 0.042773 | 31.2296 | 0.042773 | 31.2221 | 0.042773 | 31.2342 |  |  | 1.5 | 0.94748 | 1.1586 | 1.3103 | 1.3246 | 1.3463 | 1.3034 | 1.5625 | 0.98566 | 0.71854 | 1.3153 | 2.5767 | 0.21053 | 1.0875 | 0.99273 | 0.066372 | 1.4
-0.405 | 0.59489 | 0.044069 | 32.1691 | [8 empty] | 0.044069 | 32.1702 | 0.044069 | 32.162 | 0.044069 | 32.1753 |  |  | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
-0.4 | 0.42964 | 0.045365 | 33.1092 | [8 empty] | 0.045365 | 33.1104 | 0.045365 | 33.1015 | 0.045365 | 33.1159 |  |  | 0.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
-0.395 | 0.45639 | 0.046661 | 34.0488 | [8 empty] | 0.046661 | 34.0502 | 0.046661 | 34.0404 | 0.046661 | 34.0562
-0.39 | 0.48315 | 0.047957 | 34.988 | [8 empty] | 0.047957 | 34.9896 | 0.047957 | 34.9789 | 0.047957 | 34.996 |  |  | Experimental data: Zhang Guixin Plasma Soft X-ray Source for Microelectronics Lithography
-0.385 | 0.5099 | 0.049254 | 35.9268 | [8 empty] | 0.049254 | 35.9285 | 0.049254 | 35.917 | 0.049254 | 35.9355 |  |  | PhD NTU/NIE 1999 pg 206 Fig6.77b
-0.38 | 0.44066 | 0.05055 | 36.8651 | [8 empty] | 0.05055 | 36.8669 | 0.05055 | 36.8545 | 0.05055 | 36.8745 |  |  | NX2-same configuration 11.5 kV
-0.375 | 0.37141 | 0.051846 | 37.803 | [8 empty] | 0.051846 | 37.8049 | 0.051846 | 37.7915 | 0.051846 | 37.8131 |  |  | Published in IEEE Trans Plasma Sci 26 (1998) 1119
-0.37 | 0.39816 | 0.053142 | 38.7404 | [8 empty] | 0.053142 | 38.7425 | 0.053142 | 38.728 | 0.053142 | 38.7513 |  |  | Po Neon | Ysxr | Ysxr normalised
-0.365 | 0.23292 | 0.054438 | 39.6773 | [8 empty] | 0.054438 | 39.6795 | 0.054438 | 39.664 | 0.054438 | 39.689 |  |  | Torr | J | to Optimum
-0.36 | 0.25967 | 0.055734 | 40.6137 | [8 empty] | 0.055734 | 40.6161 | 0.055734 | 40.5994 | 0.055734 | 40.6263 |  |  | 1 | 2 | 0.10526
-0.355 | 0.19043 | 0.057031 | 41.5497 | [8 empty] | 0.057031 | 41.5522 | 0.057031 | 41.5343 | 0.057031 | 41.5632 |  |  | 1.5 | 9 | 0.47368
-0.35 | 0.12118 | 0.058327 | 42.4851 | [8 empty] | 0.058327 | 42.4879 | 0.058327 | 42.4687 | 0.058327 | 42.4996 |  |  | 2 | 8 | 0.42105
-0.345 | 0.14793 | 0.059623 | 43.4201 | [8 empty] | 0.059623 | 43.423 | 0.059623 | 43.4025 | 0.059623 | 43.4355 |  |  | 2.5 | 13 | 0.68421
-0.34 | 0.07869 | 0.060919 | 44.3545 | [8 empty] | 0.060919 | 44.3576 | 0.060919 | 44.3358 | 0.060919 | 44.371 |  |  | 2.8 | 15.9 | 0.83684
-0.335 | 0.10544 | 0.062215 | 45.2884 | [8 empty] | 0.062215 | 45.2918 | 0.062215 | 45.2685 | 0.062215 | 45.306 |  |  | 3 | 16.1 | 0.84737
-0.33 | 0.1322 | 0.063511 | 46.2218 | [8 empty] | 0.063511 | 46.2254 | 0.063511 | 46.2006 | 0.063511 | 46.2405 |  |  | 3.5 | 19 | 1
-0.325 | 0.06295 | 0.064807 | 47.1547 | [8 empty] | 0.064807 | 47.1585 | 0.064807 | 47.1322 | 0.064807 | 47.1745 |  |  | 4 | 14 | 0.73684
-0.32 | -0.0063 | 0.066104 | 48.087 | [8 empty] | 0.066104 | 48.091 | 0.066104 | 48.0631 | 0.066104 | 48.1081 |  |  | 4.6 | 9 | 0.47368
-0.315 | -0.07554 | 0.0674 | 49.0188 | [8 empty] | 0.0674 | 49.0231 | 0.0674 | 48.9935 | 0.0674 | 49.0411 |  |  | 5 | 2 | 0.10526
-0.31 | -0.14479 | 0.068696 | 49.95 | [8 empty] | 0.068696 | 49.9545 | 0.068696 | 49.9232 | 0.068696 | 49.9737
-0.305 | -0.21403 | 0.069992 | 50.8807 | [8 empty] | 0.069992 | 50.8855 | 0.069992 | 50.8524 | 0.069992 | 50.9057
-0.3 | -0.37928 | 0.071288 | 51.8108 | [8 empty] | 0.071288 | 51.8158 | 0.071288 | 51.7809 | 0.071288 | 51.8372
-0.295 | -0.35252 | 0.072584 | 52.7404 | [8 empty] | 0.072584 | 52.7457 | 0.072584 | 52.7087 | 0.072584 | 52.7682
-0.29 | -0.32577 | 0.07388 | 53.6693 | [8 empty] | 0.07388 | 53.6749 | 0.07388 | 53.636 | 0.07388 | 53.6987
-0.285 | -0.20302 | 0.075177 | 54.5977 | [8 empty] | 0.075177 | 54.6036 | 0.075177 | 54.5626 | 0.075177 | 54.6287
-0.28 | -0.08026 | 0.076473 | 55.5255 | [8 empty] | 0.076473 | 55.5317 | 0.076473 | 55.4885 | 0.076473 | 55.5581
-0.275 | 0.04249 | 0.077769 | 56.4526 | [8 empty] | 0.077769 | 56.4592 | 0.077769 | 56.4138 | 0.077769 | 56.4869
-0.27 | 0.06925 | 0.079065 | 57.3792 | [8 empty] | 0.079065 | 57.3861 | 0.079065 | 57.3384 | 0.079065 | 57.4152
-0.265 | 0.096 | 0.080361 | 58.3052 | [8 empty] | 0.080361 | 58.3124 | 0.080361 | 58.2623 | 0.080361 | 58.343
-0.26 | 0.21875 | 0.081657 | 59.2305 | [8 empty] | 0.081657 | 59.2381 | 0.081657 | 59.1856 | 0.081657 | 59.2702
-0.255 | 0.24551 | 0.082953 | 60.1552 | [8 empty] | 0.082953 | 60.1632 | 0.082953 | 60.1081 | 0.082953 | 60.1968
-0.25 | 0.27226 | 0.08425 | 61.0793 | [8 empty] | 0.08425 | 61.0876 | 0.08425 | 61.03 | 0.08425 | 61.1229
-0.245 | 0.39502 | 0.085546 | 62.0028 | [8 empty] | 0.085546 | 62.0114 | 0.085546 | 61.9511 | 0.085546 | 62.0483
-0.24 | 0.42177 | 0.086842 | 62.9255 | [8 empty] | 0.086842 | 62.9346 | 0.086842 | 62.8715 | 0.086842 | 62.9732
-0.235 | 0.35252 | 0.088138 | 63.8477 | [8 empty] | 0.088138 | 63.8572 | 0.088138 | 63.7912 | 0.088138 | 63.8975
-0.23 | 0.47528 | 0.089434 | 64.7692 | [8 empty] | 0.089434 | 64.7791 | 0.089434 | 64.7102 | 0.089434 | 64.8212
-0.225 | 0.40603 | 0.09073 | 65.69 | [8 empty] | 0.09073 | 65.7003 | 0.09073 | 65.6284 | 0.09073 | 65.7443
-0.22 | 0.43279 | 0.092027 | 66.6101 | [8 empty] | 0.092027 | 66.6209 | 0.092027 | 66.5459 | 0.092027 | 66.6668
-0.215 | 0.45954 | 0.093323 | 67.5296 | [8 empty] | 0.093323 | 67.5408 | 0.093323 | 67.4627 | 0.093323 | 67.5887
-0.21 | 0.3903 | 0.094619 | 68.4483 | [8 empty] | 0.094619 | 68.4601 | 0.094619 | 68.3786 | 0.094619 | 68.5099
-0.205 | 0.41705 | 0.095915 | 69.3664 | [8 empty] | 0.095915 | 69.3786 | 0.095915 | 69.2938 | 0.095915 | 69.4305
-0.2 | 0.3478 | 0.097211 | 70.2838 | [8 empty] | 0.097211 | 70.2965 | 0.097211 | 70.2083 | 0.097211 | 70.3505
-0.195 | 0.27856 | 0.098507 | 71.2005 | [8 empty] | 0.098507 | 71.2137 | 0.098507 | 71.1219 | 0.098507 | 71.2699
-0.19 | 0.20931 | 0.099803 | 72.1164 | [8 empty] | 0.099803 | 72.1302 | 0.099803 | 72.0348 | 0.099803 | 72.1886
-0.185 | 0.14007 | 0.1011 | 73.0317 | [8 empty] | 0.1011 | 73.0459 | 0.1011 | 72.9468 | 0.1011 | 73.1066
-0.18 | 0.16682 | 0.1024 | 73.9462 | [8 empty] | 0.1024 | 73.961 | 0.1024 | 73.8581 | 0.1024 | 74.024
-0.175 | 0.09757 | 0.10369 | 74.86 | [8 empty] | 0.10369 | 74.8753 | 0.10369 | 74.7685 | 0.10369 | 74.9408
-0.17 | 0.22033 | 0.10499 | 75.773 | [8 empty] | 0.10499 | 75.7889 | 0.10499 | 75.6781 | 0.10499 | 75.8569
-0.165 | 0.24708 | 0.10628 | 76.6853 | [8 empty] | 0.10628 | 76.7018 | 0.10628 | 76.5869 | 0.10628 | 76.7723
-0.16 | 0.17784 | 0.10758 | 77.5968 | [8 empty] | 0.10758 | 77.614 | 0.10758 | 77.4948 | 0.10758 | 77.687
-0.155 | 0.30059 | 0.10888 | 78.5076 | [8 empty] | 0.10888 | 78.5254 | 0.10888 | 78.4019 | 0.10888 | 78.601
-0.15 | 0.32734 | 0.11017 | 79.4176 | [8 empty] | 0.11017 | 79.436 | 0.11017 | 79.3082 | 0.11017 | 79.5144
-0.145 | 0.2581 | 0.11147 | 80.3268 | [8 empty] | 0.11147 | 80.3459 | 0.11147 | 80.2136 | 0.11147 | 80.427
-0.14 | 0.18885 | 0.11276 | 81.2353 | [8 empty] | 0.11276 | 81.255 | 0.11276 | 81.1181 | 0.11276 | 81.339
-0.135 | 0.11961 | 0.11406 | 82.143 | [8 empty] | 0.11406 | 82.1634 | 0.11406 | 82.0218 | 0.11406 | 82.2502
-0.13 | 0.14636 | 0.11536 | 83.0499 | [8 empty] | 0.11536 | 83.071 | 0.11536 | 82.9245 | 0.11536 | 83.1607
-0.125 | 0.26911 | 0.11665 | 83.956 | [8 empty] | 0.11665 | 83.9778 | 0.11665 | 83.8264 | 0.11665 | 84.0705
-0.12 | 0.39187 | 0.11795 | 84.8612 | [8 empty] | 0.11795 | 84.8838 | 0.11795 | 84.7274 | 0.11795 | 84.9796
-0.115 | 0.32262 | 0.11925 | 85.7657 | [8 empty] | 0.11925 | 85.789 | 0.11925 | 85.6275 | 0.11925 | 85.888
-0.11 | 0.25338 | 0.12054 | 86.6694 | [8 empty] | 0.12054 | 86.6934 | 0.12054 | 86.5267 | 0.12054 | 86.7956
-0.105 | 0.08813 | 0.12184 | 87.5722 | [8 empty] | 0.12184 | 87.597 | 0.12184 | 87.425 | 0.12184 | 87.7025
-0.1 | 0.11489 | 0.12313 | 88.4742 | [8 empty] | 0.12313 | 88.4998 | 0.12313 | 88.3223 | 0.12313 | 88.6086
-0.095 | 0.14164 | 0.12443 | 89.3754 | [8 empty] | 0.12443 | 89.4017 | 0.12443 | 89.2188 | 0.12443 | 89.514
-0.09 | 0.26439 | 0.12573 | 90.2757 | [8 empty] | 0.12573 | 90.3029 | 0.12573 | 90.1142 | 0.12573 | 90.4186
-0.085 | 0.38715 | 0.12702 | 91.1752 | [8 empty] | 0.12702 | 91.2032 | 0.12702 | 91.0088 | 0.12702 | 91.3225
-0.08 | 0.5099 | 0.12832 | 92.0738 | [8 empty] | 0.12832 | 92.1026 | 0.12832 | 91.9024 | 0.12832 | 92.2256
-0.075 | 0.92066 | 0.12961 | 92.9715 | [8 empty] | 0.12961 | 93.0013 | 0.12961 | 92.795 | 0.12961 | 93.1279
-0.07 | -0.20459 | 0.13091 | 93.8684 | [8 empty] | 0.13091 | 93.899 | 0.13091 | 93.6867 | 0.13091 | 94.0294
-0.065 | 1.1662 | 0.13221 | 94.7645 | [8 empty] | 0.13221 | 94.796 | 0.13221 | 94.5774 | 0.13221 | 94.9302
-0.06 | 1.3849 | 0.1335 | 95.6596 | [8 empty] | 0.1335 | 95.692 | 0.1335 | 95.4671 | 0.1335 | 95.8301
-0.055 | 6.0197 | 0.1348 | 96.5539 | [8 empty] | 0.1348 | 96.5872 | 0.1348 | 96.3558 | 0.1348 | 96.7293
-0.05 | 1.1504 | 0.1361 | 97.4472 | [8 empty] | 0.1361 | 97.4815 | 0.1361 | 97.2436 | 0.1361 | 97.6276
-0.045 | -2.9508 | 0.13739 | 98.3397 | [8 empty] | 0.13739 | 98.3749 | 0.13739 | 98.1303 | 0.13739 | 98.5252
-0.04 | -9.3561 | 0.13869 | 99.2312 | [8 empty] | 0.13869 | 99.2675 | 0.13869 | 99.016 | 0.13869 | 99.4219
-0.035 | -11.8253 | 0.13998 | 100.1219 | [8 empty] | 0.13998 | 100.1591 | 0.13998 | 99.9008 | 0.13998 | 100.3178
-0.03 | -10.8386 | 0.14128 | 101.0116 | [8 empty] | 0.14128 | 101.0499 | 0.14128 | 100.7845 | 0.14128 | 101.2129
-0.025 | -4.4758 | 0.14258 | 101.9004 | [8 empty] | 0.14258 | 101.9397 | 0.14258 | 101.6671 | 0.14258 | 102.1072
-0.02 | 1.6949 | 0.14387 | 102.7883 | [8 empty] | 0.14387 | 102.8287 | 0.14387 | 102.5488 | 0.14387 | 103.0006
-0.015 | 6.5217 | 0.14517 | 103.6753 | [8 empty] | 0.14517 | 103.7167 | 0.14517 | 103.4294 | 0.14517 | 103.8932
-0.01 | 9.6205 | 0.14646 | 104.5613 | [8 empty] | 0.14646 | 104.6038 | 0.14646 | 104.309 | 0.14646 | 104.785
-0.005 | 10.5112 | 0.14776 | 105.4463 | [8 empty] | 0.14776 | 105.49 | 0.14776 | 105.1875 | 0.14776 | 105.6759
6.62e-18 | 12.458 | 0.14906 | 106.3305 | [8 empty] | 0.14906 | 106.3752 | 0.14906 | 106.0649 | 0.14906 | 106.5659
0.005 | 13.7327 | 0.15035 | 107.2136 | [8 empty] | 0.15035 | 107.2595 | 0.15035 | 106.9413 | 0.15035 | 107.4551
0.01 | 16.8315 | 0.15165 | 108.0958 | [8 empty] | 0.15165 | 108.1429 | 0.15165 | 107.8166 | 0.15165 | 108.3434
0.015 | 16.4742 | 0.15295 | 108.977 | [8 empty] | 0.15295 | 109.0253 | 0.15295 | 108.6909 | 0.15295 | 109.2309
0.02 | 15.349 | 0.15424 | 109.8573 | [8 empty] | 0.15424 | 109.9067 | 0.15424 | 109.564 | 0.15424 | 110.1174
0.025 | 15.6637 | 0.15554 | 110.7365 | [8 empty] | 0.15554 | 110.7872 | 0.15554 | 110.4361 | 0.15554 | 111.0031
0.03 | 19.5305 | 0.15683 | 111.6148 | [8 empty] | 0.15683 | 111.6667 | 0.15683 | 111.307 | 0.15683 | 111.8879
0.035 | 25.7972 | 0.15813 | 112.4921 | [8 empty] | 0.15813 | 112.5452 | 0.15813 | 112.1769 | 0.15813 | 112.7718
0.04 | 31.296 | 0.15943 | 113.3684 | [8 empty] | 0.15943 | 113.4228 | 0.15943 | 113.0456 | 0.15943 | 113.6548
0.045 | 37.0827 | 0.16072 | 114.2437 | [8 empty] | 0.16072 | 114.2994 | 0.16072 | 113.9133 | 0.16072 | 114.5369
0.05 | 39.2215 | 0.16202 | 115.1179 | [8 empty] | 0.16202 | 115.175 | 0.16202 | 114.7798 | 0.16202 | 115.4181
0.055 | 40.4003 | 0.16331 | 115.9912 | [8 empty] | 0.16331 | 116.0495 | 0.16331 | 115.6452 | 0.16331 | 116.2984
0.06 | 39.083 | 0.16461 | 116.8634 | [8 empty] | 0.16461 | 116.9231 | 0.16461 | 116.5094 | 0.16461 | 117.1778
0.065 | 41.1258 | 0.16591 | 117.7346 | [8 empty] | 0.16591 | 117.7957 | 0.16591 | 117.3725 | 0.16591 | 118.0562
0.07 | 43.7445 | 0.1672 | 118.6048 | [8 empty] | 0.1672 | 118.6673 | 0.1672 | 118.2344 | 0.1672 | 118.9337
0.075 | 47.1313 | 0.1685 | 119.4739 | [8 empty] | 0.1685 | 119.5378 | 0.1685 | 119.0952 | 0.1685 | 119.8103
0.08 | 49.174 | 0.1698 | 120.342 | [8 empty] | 0.1698 | 120.4073 | 0.1698 | 119.9549 | 0.1698 | 120.686
0.085 | 53.8088 | 0.17109 | 121.209 | [8 empty] | 0.17109 | 121.2758 | 0.17109 | 120.8133 | 0.17109 | 121.5606
0.09 | 58.7315 | 0.17239 | 122.075 | [8 empty] | 0.17239 | 122.1433 | 0.17239 | 121.6706 | 0.17239 | 122.4344
0.095 | 62.1183 | 0.17368 | 122.9399 | [8 empty] | 0.17368 | 123.0097 | 0.17368 | 122.5268 | 0.17368 | 123.3072
0.1 | 65.409 | 0.17498 | 123.8038 | [8 empty] | 0.17498 | 123.875 | 0.17498 | 123.3817 | 0.17498 | 124.179
0.105 | 68.2198 | 0.17628 | 124.6666 | [8 empty] | 0.17628 | 124.7393 | 0.17628 | 124.2354 | 0.17628 | 125.0499
0.11 | 71.1266 | 0.17757 | 125.5283 | [8 empty] | 0.17757 | 125.6026 | 0.17757 | 125.088 | 0.17757 | 125.9198
0.115 | 71.7293 | 0.17887 | 126.3889 | [8 empty] | 0.17887 | 126.4648 | 0.17887 | 125.9393 | 0.17887 | 126.7887
0.12 | 75.2121 | 0.18016 | 127.2484 | [8 empty] | 0.18016 | 127.3259 | 0.18016 | 126.7895 | 0.18016 | 127.6566
0.125 | 78.5028 | 0.18146 | 128.1068 | [8 empty] | 0.18146 | 128.186 | 0.18146 | 127.6384 | 0.18146 | 128.5236
0.13 | 82.0816 | 0.18276 | 128.9642 | [8 empty] | 0.18276 | 129.0449 | 0.18276 | 128.4861 | 0.18276 | 129.3896
0.135 | 83.6443 | 0.18405 | 129.8204 | [8 empty] | 0.18405 | 129.9028 | 0.18405 | 129.3325 | 0.18405 | 130.2545
0.14 | 87.4151 | 0.18535 | 130.6755 | [8 empty] | 0.18535 | 130.7596 | 0.18535 | 130.1778 | 0.18535 | 131.1185
0.145 | 91.1858 | 0.18665 | 131.5295 | [8 empty] | 0.18665 | 131.6153 | 0.18665 | 131.0218 | 0.18665 | 131.9815
0.15 | 94.7646 | 0.18794 | 132.3824 | [8 empty] | 0.18794 | 132.4699 | 0.18794 | 131.8645 | 0.18794 | 132.8434
0.155 | 97.3833 | 0.18924 | 133.2342 | [8 empty] | 0.18924 | 133.3234 | 0.18924 | 132.706 | 0.18924 | 133.7044
0.16 | 101.1541 | 0.19053 | 134.0848 | [8 empty] | 0.19053 | 134.1758 | 0.19053 | 133.5463 | 0.19053 | 134.5643
0.165 | 105.1169 | 0.19183 | 134.9343 | [8 empty] | 0.19183 | 135.0271 | 0.19183 | 134.3853 | 0.19183 | 135.4232
0.17 | 107.5436 | 0.19313 | 135.7826 | [8 empty] | 0.19313 | 135.8772 | 0.19313 | 135.223 | 0.19313 | 136.2811
0.175 | 110.2584 | 0.19442 | 136.6298 | [8 empty] | 0.19442 | 136.7262 | 0.19442 | 136.0595 | 0.19442 | 137.1379
0.18 | 113.2611 | 0.19572 | 137.4759 | [8 empty] | 0.19572 | 137.5741 | 0.19572 | 136.8946 | 0.19572 | 137.9937
0.185 | 116.8399 | 0.19701 | 138.3208 | [8 empty] | 0.19701 | 138.4209 | 0.19701 | 137.7285 | 0.19701 | 138.8485
0.19 | 119.5546 | 0.19831 | 139.1645 | [8 empty] | 0.19831 | 139.2665 | 0.19831 | 138.5611 | 0.19831 | 139.7022
0.195 | 122.7494 | 0.19961 | 140.007 | [8 empty] | 0.19961 | 140.111 | 0.19961 | 139.3924 | 0.19961 | 140.5549
0.2 | 126.5201 | 0.2009 | 140.8484 | [8 empty] | 0.2009 | 140.9543 | 0.2009 | 140.2224 | 0.2009 | 141.4065
0.205 | 130.0989 | 0.2022 | 141.6886 | [8 empty] | 0.2022 | 141.7964 | 0.2022 | 141.0511 | 0.2022 | 142.257
0.21 | 133.3896 | 0.2035 | 142.5276 | [8 empty] | 0.2035 | 142.6374 | 0.2035 | 141.8785 | 0.2035 | 143.1065
0.215 | 136.4884 | 0.20479 | 143.3655 | [8 empty] | 0.20479 | 143.4773 | 0.20479 | 142.7046 | 0.20479 | 143.9549
0.22 | 139.6831 | 0.20609 | 144.2021 | [8 empty] | 0.20609 | 144.3159 | 0.20609 | 143.5294 | 0.20609 | 144.8022
0.225 | 142.1099 | 0.20738 | 145.0376 | [8 empty] | 0.20738 | 145.1534 | 0.20738 | 144.3528 | 0.20738 | 145.6484
0.23 | 145.0167 | 0.20868 | 145.8718 | [8 empty] | 0.20868 | 145.9897 | 0.20868 | 145.1749 | 0.20868 | 146.4936
0.235 | 148.5954 | 0.20998 | 146.7048 | [8 empty] | 0.20998 | 146.8248 | 0.20998 | 145.9956 | 0.20998 | 147.3377
0.24 | 152.1742 | 0.21127 | 147.5366 | [8 empty] | 0.21127 | 147.6587 | 0.21127 | 146.8151 | 0.21127 | 148.1807
0.245 | 154.9849 | 0.21257 | 148.3672 | [8 empty] | 0.21257 | 148.4915 | 0.21257 | 147.6331 | 0.21257 | 149.0225
0.25 | 157.9877 | 0.21386 | 149.1966 | [8 empty] | 0.21386 | 149.323 | 0.21386 | 148.4498 | 0.21386 | 149.8633
0.255 | 161.4704 | 0.21516 | 150.0247 | [8 empty] | 0.21516 | 150.1533 | 0.21516 | 149.2652 | 0.21516 | 150.703
0.26 | 164.6652 | 0.21646 | 150.8517 | [8 empty] | 0.21646 | 150.9824 | 0.21646 | 150.0792 | 0.21646 | 151.5415
0.265 | 167.4759 | 0.21775 | 151.6773 | [8 empty] | 0.21775 | 151.8103 | 0.21775 | 150.8918 | 0.21775 | 152.379
0.27 | 170.3827 | 0.21905 | 152.5018 | [8 empty] | 0.21905 | 152.637 | 0.21905 | 151.7031 | 0.21905 | 153.2153
0.275 | 173.5774 | 0.22035 | 153.3249 | [8 empty] | 0.22035 | 153.4624 | 0.22035 | 152.5129 | 0.22035 | 154.0505
0.28 | 176.1962 | 0.22164 | 154.1469 | [8 empty] | 0.22164 | 154.2866 | 0.22164 | 153.3214 | 0.22164 | 154.8845
0.285 | 178.719 | 0.22294 | 154.9675 | [8 empty] | 0.22294 | 155.1096 | 0.22294 | 154.1285 | 0.22294 | 155.7174
0.29 | 181.9137 | 0.22423 | 155.787 | [8 empty] | 0.22423 | 155.9314 | 0.22423 | 154.9342 | 0.22423 | 156.5492
0.295 | 185.2045 | 0.22553 | 156.6051 | [8 empty] | 0.22553 | 156.7519 | 0.22553 | 155.7386 | 0.22553 | 157.3799
0.3 | 188.1112 | 0.22683 | 157.422 | [8 empty] | 0.22683 | 157.5711 | 0.22683 | 156.5415 | 0.22683 | 158.2093
0.305 | 191.018 | 0.22812 | 158.2376 | [8 empty] | 0.22812 | 158.3891 | 0.22812 | 157.343 | 0.22812 | 159.0377
0.31 | 194.5007 | 0.22942 | 159.0519 | [8 empty] | 0.22942 | 159.2059 | 0.22942 | 158.1431 | 0.22942 | 159.8649
0.315 | 197.6955 | 0.23071 | 159.8649 | [8 empty] | 0.23071 | 160.0214 | 0.23071 | 158.9417 | 0.23071 | 160.6909
0.32 | 200.2182 | 0.23201 | 160.6767 | [8 empty] | 0.23201 | 160.8356 | 0.23201 | 159.739 | 0.23201 | 161.5157
0.325 | 202.933 | 0.23331 | 161.4871 | [8 empty] | 0.23331 | 161.6485 | 0.23331 | 160.5348 | 0.23331 | 162.3394
0.33 | 205.9357 | 0.2346 | 162.2963 | [8 empty] | 0.2346 | 162.4602 | 0.2346 | 161.3292 | 0.2346 | 163.1619
0.335 | 208.6505 | 0.2359 | 163.1041 | [8 empty] | 0.2359 | 163.2706 | 0.2359 | 162.1222 | 0.2359 | 163.9832
0.34 | 211.2692 | 0.2372 | 163.9107 | [8 empty] | 0.2372 | 164.0797 | 0.2372 | 162.9137 | 0.2372 | 164.8034
0.345 | 214.176 | 0.23849 | 164.7159 | [8 empty] | 0.23849 | 164.8875 | 0.23849 | 163.7037 | 0.23849 | 165.6223
0.35 | 217.5628 | 0.23979 | 165.5198 | [8 empty] | 0.23979 | 165.694 | 0.23979 | 164.4924 | 0.23979 | 166.4401
0.355 | 220.4695 | 0.24108 | 166.3224 | [8 empty] | 0.24108 | 166.4993 | 0.24108 | 165.2795 | 0.24108 | 167.2567
0.36 | 223.0883 | 0.24238 | 167.1237 | [8 empty] | 0.24238 | 167.3032 | 0.24238 | 166.0652 | 0.24238 | 168.072
0.365 | 225.995 | 0.24368 | 167.9236 | [8 empty] | 0.24368 | 168.1058 | 0.24368 | 166.8495 | 0.24368 | 168.8862
0.37 | 228.9978 | 0.24497 | 168.7222 | [8 empty] | 0.24497 | 168.9071 | 0.24497 | 167.6323 | 0.24497 | 169.6992
0.375 | 231.5205 | 0.24627 | 169.5195 | [8 empty] | 0.24627 | 169.7071 | 0.24627 | 168.4136 | 0.24627 | 170.5109
0.38 | 234.0433 | 0.24756 | 170.3154 | [8 empty] | 0.24756 | 170.5058 | 0.24756 | 169.1934 | 0.24756 | 171.3214
0.385 | 236.854 | 0.24886 | 171.11 | [8 empty] | 0.24886 | 171.3031 | 0.24886 | 169.9717 | 0.24886 | 172.1307
0.39 | 239.2808 | 0.25016 | 171.9032 | [8 empty] | 0.25016 | 172.0992 | 0.25016 | 170.7486 | 0.25016 | 172.9388
0.395 | 241.6115 | 0.25145 | 172.6951 | [8 empty] | 0.25145 | 172.8939 | 0.25145 | 171.524 | 0.25145 | 173.7457
0.4 | 244.3263 | 0.25275 | 173.4856 | [8 empty] | 0.25275 | 173.6872 | 0.25275 | 172.2978 | 0.25275 | 174.5513
0.405 | 247.329 | 0.25405 | 174.2747 | [8 empty] | 0.25405 | 174.4792 | 0.25405 | 173.0702 | 0.25405 | 175.3557
0.41 | 250.1398 | 0.25534 | 175.0625 | [8 empty] | 0.25534 | 175.2699 | 0.25534 | 173.8411 | 0.25534 | 176.1588
0.415 | 252.8546 | 0.25664 | 175.8489 | [8 empty] | 0.25664 | 176.0592 | 0.25664 | 174.6105 | 0.25664 | 176.9607
0.42 | 255.4733 | 0.25793 | 176.6339 | [8 empty] | 0.25793 | 176.8472 | 0.25793 | 175.3783 | 0.25793 | 177.7613
0.425 | 258.0921 | 0.25923 | 177.4176 | [8 empty] | 0.25923 | 177.6338 | 0.25923 | 176.1447 | 0.25923 | 178.5607
0.43 | 260.3268 | 0.26053 | 178.1999 | [8 empty] | 0.26053 | 178.4191 | 0.26053 | 176.9095 | 0.26053 | 179.3589
0.435 | 262.7536 | 0.26182 | 178.9807 | [8 empty] | 0.26182 | 179.2029 | 0.26182 | 177.6728 | 0.26182 | 180.1557
0.44 | 265.6603 | 0.26312 | 179.7602 | [8 empty] | 0.26312 | 179.9855 | 0.26312 | 178.4346 | 0.26312 | 180.9514
0.445 | 268.2791 | 0.26441 | 180.5383 | [8 empty] | 0.26441 | 180.7666 | 0.26441 | 179.1949 | 0.26441 | 181.7457
0.45 | 270.4178 | 0.26571 | 181.315 | [8 empty] | 0.26571 | 181.5464 | 0.26571 | 179.9536 | 0.26571 | 182.5388
0.455 | 272.5566 | 0.26701 | 182.0903 | [8 empty] | 0.26701 | 182.3248 | 0.26701 | 180.7108 | 0.26701 | 183.3306
0.46 | 274.8873 | 0.2683 | 182.8642 | [8 empty] | 0.2683 | 183.1018 | 0.2683 | 181.4664 | 0.2683 | 184.1211
0.465 | 277.2181 | 0.2696 | 183.6366 | [8 empty] | 0.2696 | 183.8774 | 0.2696 | 182.2205 | 0.2696 | 184.9103
0.47 | 279.4529 | 0.27089 | 184.4077 | [8 empty] | 0.27089 | 184.6516 | 0.27089 | 182.973 | 0.27089 | 185.6982
0.475 | 281.7836 | 0.27219 | 185.1773 | [8 empty] | 0.27219 | 185.4244 | 0.27219 | 183.724 | 0.27219 | 186.4849
0.48 | 284.3064 | 0.27349 | 185.9455 | [8 empty] | 0.27349 | 186.1959 | 0.27349 | 184.4735 | 0.27349 | 187.2702
0.485 | 286.4451 | 0.27478 | 186.7123 | [8 empty] | 0.27478 | 186.9659 | 0.27478 | 185.2213 | 0.27478 | 188.0543
0.49 | 288.4879 | 0.27608 | 187.4777 | [8 empty] | 0.27608 | 187.7345 | 0.27608 | 185.9676 | 0.27608 | 188.837
0.495 | 290.6266 | 0.27738 | 188.2416 | [8 empty] | 0.27738 | 188.5017 | 0.27738 | 186.7124 | 0.27738 | 189.6185
0.5 | 293.0534 | 0.27867 | 189.004 | [8 empty] | 0.27867 | 189.2675 | 0.27867 | 187.4555 | 0.27867 | 190.3986
0.505 | 295.0961 | 0.27997 | 189.7651 | [8 empty] | 0.27997 | 190.0319 | 0.27997 | 188.1971 | 0.27997 | 191.1774
0.51 | 297.0429 | 0.28126 | 190.5247 | [8 empty] | 0.28126 | 190.7948 | 0.28126 | 188.9371 | 0.28126 | 191.9549
0.515 | 298.8936 | 0.28256 | 191.2828 | [8 empty] | 0.28256 | 191.5564 | 0.28256 | 189.6756 | 0.28256 | 192.7311
0.52 | 300.7444 | 0.28386 | 192.0395 | [8 empty] | 0.28386 | 192.3165 | 0.28386 | 190.4124 | 0.28386 | 193.5059
0.525 | 302.4991 | 0.28515 | 192.7947 | [8 empty] | 0.28515 | 193.0751 | 0.28515 | 191.1476 | 0.28515 | 194.2795
0.53 | 304.8299 | 0.28645 | 193.5485 | [8 empty] | 0.28645 | 193.8324 | 0.28645 | 191.8813 | 0.28645 | 195.0516
0.535 | 307.1607 | 0.28774 | 194.3008 | [8 empty] | 0.28774 | 194.5881 | 0.28774 | 192.6134 | 0.28774 | 195.8225
0.54 | 309.0114 | 0.28904 | 195.0516 | [8 empty] | 0.28904 | 195.3425 | 0.28904 | 193.3438 | 0.28904 | 196.592
0.545 | 310.7662 | 0.29034 | 195.801 | [8 empty] | 0.29034 | 196.0954 | 0.29034 | 194.0727 | 0.29034 | 197.3602
0.55 | 312.6169 | 0.29163 | 196.5489 | [8 empty] | 0.29163 | 196.8468 | 0.29163 | 194.7999 | 0.29163 | 198.127
0.555 | 314.3717 | 0.29293 | 197.2953 | [8 empty] | 0.29293 | 197.5968 | 0.29293 | 195.5256 | 0.29293 | 198.8925
0.56 | 316.0304 | 0.29423 | 198.0402 | [8 empty] | 0.29423 | 198.3454 | 0.29423 | 196.2496 | 0.29423 | 199.6566
0.565 | 317.6892 | 0.29552 | 198.7837 | [8 empty] | 0.29552 | 199.0924 | 0.29552 | 196.972 | 0.29552 | 200.4194
0.57 | 319.4439 | 0.29682 | 199.5257 | [8 empty] | 0.29682 | 199.838 | 0.29682 | 197.6928 | 0.29682 | 201.1808
0.575 | 320.8147 | 0.29811 | 200.2661 | [8 empty] | 0.29811 | 200.5822 | 0.29811 | 198.412 | 0.29811 | 201.9408
0.58 | 322.3774 | 0.29941 | 201.0051 | [8 empty] | 0.29941 | 201.3248 | 0.29941 | 199.1296 | 0.29941 | 202.6995
0.585 | 324.0362 | 0.30071 | 201.7426 | [8 empty] | 0.30071 | 202.066 | 0.30071 | 199.8455 | 0.30071 | 203.4568
0.59 | 325.599 | 0.302 | 202.4786 | [8 empty] | 0.302 | 202.8058 | 0.302 | 200.5598 | 0.302 | 204.2127
0.595 | 327.0657 | 0.3033 | 203.213 | [8 empty] | 0.3033 | 203.544 | 0.3033 | 201.2724 | 0.3033 | 204.9673
0.6 | 328.1485 | 0.30459 | 203.946 | [8 empty] | 0.30459 | 204.2807 | 0.30459 | 201.9834 | 0.30459 | 205.7204
0.605 | 329.7112 | 0.30589 | 204.6775 | [8 empty] | 0.30589 | 205.016 | 0.30589 | 202.6928 | 0.30589 | 206.4722
0.61 | 331.178 | 0.30719 | 205.4074 | [8 empty] | 0.30719 | 205.7498 | 0.30719 | 203.4005 | 0.30719 | 207.2226
0.615 | 332.6447 | 0.30848 | 206.1358 | [8 empty] | 0.30848 | 206.4821 | 0.30848 | 204.1066 | 0.30848 | 207.9716
0.62 | 334.2075 | 0.30978 | 206.8627 | [8 empty] | 0.30978 | 207.2128 | 0.30978 | 204.811 | 0.30978 | 208.7192
0.625 | 335.6742 | 0.31108 | 207.5881 | [8 empty] | 0.31108 | 207.9421 | 0.31108 | 205.5138 | 0.31108 | 209.4654
0.63 | 337.141 | 0.31237 | 208.312 | [8 empty] | 0.31237 | 208.6699 | 0.31237 | 206.215 | 0.31237 | 210.2102
0.635 | 338.4157 | 0.31367 | 209.0343 | [8 empty] | 0.31367 | 209.3962 | 0.31367 | 206.9144 | 0.31367 | 210.9537
0.64 | 339.4985 | 0.31496 | 209.7551 | [8 empty] | 0.31496 | 210.1209 | 0.31496 | 207.6122 | 0.31496 | 211.6957
0.645 | 340.8692 | 0.31626 | 210.4743 | [8 empty] | 0.31626 | 210.8442 | 0.31626 | 208.3084 | 0.31626 | 212.4362
0.65 | 341.856 | 0.31756 | 211.192 | [8 empty] | 0.31756 | 211.5659 | 0.31756 | 209.0028 | 0.31756 | 213.1754
0.655 | 342.8428 | 0.31885 | 211.9082 | [8 empty] | 0.31885 | 212.2861 | 0.31885 | 209.6957 | 0.31885 | 213.9132
0.66 | 343.8295 | 0.32015 | 212.6228 | [8 empty] | 0.32015 | 213.0048 | 0.32015 | 210.3868 | 0.32015 | 214.6495
0.665 | 345.0083 | 0.32144 | 213.3359 | [8 empty] | 0.32144 | 213.722 | 0.32144 | 211.0763 | 0.32144 | 215.3844
0.67 | 346.187 | 0.32274 | 214.0475 | [8 empty] | 0.32274 | 214.4376 | 0.32274 | 211.764 | 0.32274 | 216.1179
0.675 | 347.3658 | 0.32404 | 214.7574 | [8 empty] | 0.32404 | 215.1517 | 0.32404 | 212.4501 | 0.32404 | 216.85
0.68 | 348.5445 | 0.32533 | 215.4659 | [8 empty] | 0.32533 | 215.8643 | 0.32533 | 213.1346 | 0.32533 | 217.5806
0.685 | 349.6273 | 0.32663 | 216.1727 | [8 empty] | 0.32663 | 216.5753 | 0.32663 | 213.8173 | 0.32663 | 218.3098
0.69 | 350.422 | 0.32793 | 216.878 | [8 empty] | 0.32793 | 217.2848 | 0.32793 | 214.4984 | 0.32793 | 219.0376
0.695 | 351.4088 | 0.32922 | 217.5818 | [8 empty] | 0.32922 | 217.9928 | 0.32922 | 215.1777 | 0.32922 | 219.7639
0.7 | 352.4915 | 0.33052 | 218.2839 | [8 empty] | 0.33052 | 218.6992 | 0.33052 | 215.8554 | 0.33052 | 220.4888
0.705 | 353.5743 | 0.33181 | 218.9845 | [8 empty] | 0.33181 | 219.4041 | 0.33181 | 216.5314 | 0.33181 | 221.2122
0.71 | 354.369 | 0.33311 | 219.6835 | [8 empty] | 0.33311 | 220.1074 | 0.33311 | 217.2057 | 0.33311 | 221.9342
0.715 | 355.1638 | 0.33441 | 220.381 | [8 empty] | 0.33441 | 220.8091 | 0.33441 | 217.8783 | 0.33441 | 222.6547
0.72 | 355.9586 | 0.3357 | 221.0769 | [8 empty] | 0.3357 | 221.5093 | 0.3357 | 218.5491 | 0.3357 | 223.3738
0.725 | 356.3693 | 0.337 | 221.7712 | [8 empty] | 0.337 | 222.208 | 0.337 | 219.2183 | 0.337 | 224.0914
0.73 | 357.2601 | 0.33829 | 222.4639 | [8 empty] | 0.33829 | 222.9051 | 0.33829 | 219.8858 | 0.33829 | 224.8076
0.735 | 358.0548 | 0.33959 | 223.155 | [8 empty] | 0.33959 | 223.6006 | 0.33959 | 220.5516 | 0.33959 | 225.5222
0.74 | 358.7536 | 0.34089 | 223.8445 | [8 empty] | 0.34089 | 224.2945 | 0.34089 | 221.2156 | 0.34089 | 226.2355
0.745 | 359.4523 | 0.34218 | 224.5325 | [8 empty] | 0.34218 | 224.9869 | 0.34218 | 221.878 | 0.34218 | 226.9472
0.75 | 359.9591 | 0.34348 | 225.2188 | [8 empty] | 0.34348 | 225.6777 | 0.34348 | 222.5386 | 0.34348 | 227.6575
0.755 | 360.6578 | 0.34478 | 225.9036 | [8 empty] | 0.34478 | 226.367 | 0.34478 | 223.1975 | 0.34478 | 228.3663
0.76 | 361.4526 | 0.34607 | 226.5867 | [8 empty] | 0.34607 | 227.0546 | 0.34607 | 223.8547 | 0.34607 | 229.0736
0.765 | 361.8633 | 0.34737 | 227.2683 | [8 empty] | 0.34737 | 227.7407 | 0.34737 | 224.5102 | 0.34737 | 229.7794
0.77 | 362.4661 | 0.34866 | 227.9482 | [8 empty] | 0.34866 | 228.4252 | 0.34866 | 225.1639 | 0.34866 | 230.4838
0.775 | 362.9729 | 0.34996 | 228.6266 | [8 empty] | 0.34996 | 229.1081 | 0.34996 | 225.816 | 0.34996 | 231.1867
0.78 | 363.4796 | 0.35126 | 229.3033 | [8 empty] | 0.35126 | 229.7895 | 0.35126 | 226.4663 | 0.35126 | 231.8881
0.785 | 363.9864 | 0.35255 | 229.9784 | [8 empty] | 0.35255 | 230.4692 | 0.35255 | 227.1149 | 0.35255 | 232.5879
0.79 | 364.3971 | 0.35385 | 230.6519 | [8 empty] | 0.35385 | 231.1474 | 0.35385 | 227.7617 | 0.35385 | 233.2863
0.795 | 364.8079 | 0.35514 | 231.3238 | [8 empty] | 0.35514 | 231.8239 | 0.35514 | 228.4068 | 0.35514 | 233.9832
0.8 | 365.3146 | 0.35644 | 231.9941 | [8 empty] | 0.35644 | 232.4989 | 0.35644 | 229.0502 | 0.35644 | 234.6786
0.805 | 365.6294 | 0.35774 | 232.6628 | [8 empty] | 0.35774 | 233.1722 | 0.35774 | 229.6918 | 0.35774 | 235.3725
0.81 | 366.0401 | 0.35903 | 233.3298 | [8 empty] | 0.35903 | 233.844 | 0.35903 | 230.3318 | 0.35903 | 236.0649
0.815 | 366.4509 | 0.36033 | 233.9952 | [8 empty] | 0.36033 | 234.5141 | 0.36033 | 230.9699 | 0.36033 | 236.7558
0.82 | 366.6696 | 0.36163 | 234.659 | [8 empty] | 0.36163 | 235.1827 | 0.36163 | 231.6064 | 0.36163 | 237.4452
0.825 | 367.0804 | 0.36292 | 235.3212 | [8 empty] | 0.36292 | 235.8496 | 0.36292 | 232.241 | 0.36292 | 238.1331
0.83 | 367.2991 | 0.36422 | 235.9817 | [8 empty] | 0.36422 | 236.515 | 0.36422 | 232.874 | 0.36422 | 238.8194
0.835 | 367.7099 | 0.36551 | 236.6406 | [8 empty] | 0.36551 | 237.1787 | 0.36551 | 233.5052 | 0.36551 | 239.5042
0.84 | 367.9287 | 0.36681 | 237.2978 | [8 empty] | 0.36681 | 237.8408 | 0.36681 | 234.1346 | 0.36681 | 240.1875
0.845 | 368.2434 | 0.36811 | 237.9535 | [8 empty] | 0.36811 | 238.5013 | 0.36811 | 234.7623 | 0.36811 | 240.8693
0.85 | 368.3662 | 0.3694 | 238.6074 | [8 empty] | 0.3694 | 239.1601 | 0.3694 | 235.3883 | 0.3694 | 241.5496
0.855 | 368.4889 | 0.3707 | 239.2598 | [8 empty] | 0.3707 | 239.8174 | 0.3707 | 236.0125 | 0.3707 | 242.2283
0.86 | 368.7077 | 0.37199 | 239.9105 | [8 empty] | 0.37199 | 240.473 | 0.37199 | 236.6349 | 0.37199 | 242.9055
0.865 | 368.8304 | 0.37329 | 240.5595 | [8 empty] | 0.37329 | 241.127 | 0.37329 | 237.2556 | 0.37329 | 243.5812
0.87 | 368.9532 | 0.37459 | 241.2069 | [8 empty] | 0.37459 | 241.7794 | 0.37459 | 237.8746 | 0.37459 | 244.2554
0.875 | 369.0759 | 0.37588 | 241.8527 | [8 empty] | 0.37588 | 242.4301 | 0.37588 | 238.4917 | 0.37588 | 244.928
0.88 | 369.1987 | 0.37718 | 242.4968 | [8 empty] | 0.37718 | 243.0793 | 0.37718 | 239.1072 | 0.37718 | 245.599
0.885 | 369.3214 | 0.37848 | 243.1393 | [8 empty] | 0.37848 | 243.7267 | 0.37848 | 239.7208 | 0.37848 | 246.2686
0.89 | 369.2522 | 0.37977 | 243.78 | [8 empty] | 0.37977 | 244.3726 | 0.37977 | 240.3327 | 0.37977 | 246.9365
0.895 | 369.279 | 0.38107 | 244.4192 | [8 empty] | 0.38107 | 245.0168 | 0.38107 | 240.9429 | 0.38107 | 247.603
0.9 | 369.3057 | 0.38236 | 245.0567 | [8 empty] | 0.38236 | 245.6594 | 0.38236 | 241.5512 | 0.38236 | 248.2679
0.905 | 369.3325 | 0.38366 | 245.6925 | [8 empty] | 0.38366 | 246.3003 | 0.38366 | 242.1578 | 0.38366 | 248.9312
0.91 | 369.2632 | 0.38496 | 246.3267 | [8 empty] | 0.38496 | 246.9396 | 0.38496 | 242.7627 | 0.38496 | 249.593
0.915 | 369.194 | 0.38625 | 246.9591 | [8 empty] | 0.38625 | 247.5772 | 0.38625 | 243.3657 | 0.38625 | 250.2532
0.92 | 369.1247 | 0.38755 | 247.59 | [8 empty] | 0.38755 | 248.2132 | 0.38755 | 243.967 | 0.38755 | 250.9119
0.925 | 369.0555 | 0.38884 | 248.2191 | [8 empty] | 0.38884 | 248.8476 | 0.38884 | 244.5666 | 0.38884 | 251.569
0.93 | 369.0822 | 0.39014 | 248.8466 | [8 empty] | 0.39014 | 249.4803 | 0.39014 | 245.1643 | 0.39014 | 252.2246
0.935 | 369.205 | 0.39144 | 249.4725 | [8 empty] | 0.39144 | 250.1113 | 0.39144 | 245.7603 | 0.39144 | 252.8786
0.94 | 369.2317 | 0.39273 | 250.0966 | [8 empty] | 0.39273 | 250.7407 | 0.39273 | 246.3545 | 0.39273 | 253.531
0.945 | 369.1625 | 0.39403 | 250.7191 | [8 empty] | 0.39403 | 251.3685 | 0.39403 | 246.947 | 0.39403 | 254.1819
0.95 | 368.9012 | 0.39533 | 251.3399 | [8 empty] | 0.39533 | 251.9946 | 0.39533 | 247.5377 | 0.39533 | 254.8311
0.955 | 368.448 | 0.39662 | 251.959 | [8 empty] | 0.39662 | 252.619 | 0.39662 | 248.1265 | 0.39662 | 255.4789
0.96 | 368.4748 | 0.39792 | 252.5765 | [8 empty] | 0.39792 | 253.2417 | 0.39792 | 248.7137 | 0.39792 | 256.125
0.965 | 368.4055 | 0.39921 | 253.1922 | [8 empty] | 0.39921 | 253.8629 | 0.39921 | 249.299 | 0.39921 | 256.7696
0.97 | 368.3363 | 0.40051 | 253.8063 | [8 empty] | 0.40051 | 254.4823 | 0.40051 | 249.8825 | 0.40051 | 257.4126
0.975 | 368.267 | 0.40181 | 254.4187 | [8 empty] | 0.40181 | 255.1001 | 0.40181 | 250.4643 | 0.40181 | 258.054
0.98 | 368.1978 | 0.4031 | 255.0294 | [8 empty] | 0.4031 | 255.7162 | 0.4031 | 251.0443 | 0.4031 | 258.6938
0.985 | 367.8405 | 0.4044 | 255.6385 | [8 empty] | 0.4044 | 256.3306 | 0.4044 | 251.6225 | 0.4044 | 259.3321
0.99 | 367.5793 | 0.40569 | 256.2458 | [8 empty] | 0.40569 | 256.9434 | 0.40569 | 252.1989 | 0.40569 | 259.9687
0.995 | 367.222 | 0.40699 | 256.8514 | [8 empty] | 0.40699 | 257.5545 | 0.40699 | 252.7736 | 0.40699 | 260.6038
1 | 366.7688 | 0.40829 | 257.4554 | [8 empty] | 0.40829 | 258.1639 | 0.40829 | 253.3465 | 0.40829 | 261.2373
1.005 | 366.3155 | 0.40958 | 258.0577 | [8 empty] | 0.40958 | 258.7717 | 0.40958 | 253.9175 | 0.40958 | 261.8692
1.01 | 365.9583 | 0.41088 | 258.6583 | [8 empty] | 0.41088 | 259.3778 | 0.41088 | 254.4868 | 0.41088 | 262.4995
1.015 | 365.601 | 0.41218 | 259.2571 | [8 empty] | 0.41218 | 259.9822 | 0.41218 | 255.0543 | 0.41218 | 263.1282
1.02 | 365.0518 | 0.41347 | 259.8543 | [8 empty] | 0.41347 | 260.5849 | 0.41347 | 255.62 | 0.41347 | 263.7553
1.025 | 364.5026 | 0.41477 | 260.4498 | [8 empty] | 0.41477 | 261.1859 | 0.41477 | 256.184 | 0.41477 | 264.3808
1.03 | 364.1453 | 0.41606 | 261.0436 | [8 empty] | 0.41606 | 261.7853 | 0.41606 | 256.7461 | 0.41606 | 265.0047
1.035 | 363.7881 | 0.41736 | 261.6357 | [8 empty] | 0.41736 | 262.383 | 0.41736 | 257.3065 | 0.41736 | 265.6271
1.04 | 363.3348 | 0.41866 | 262.2261 | [8 empty] | 0.41866 | 262.979 | 0.41866 | 257.865 | 0.41866 | 266.2478
1.045 | 362.7856 | 0.41995 | 262.8148 | [8 empty] | 0.41995 | 263.5733 | 0.41995 | 258.4218 | 0.41995 | 266.8669
1.05 | 362.4283 | 0.42125 | 263.4018 | [8 empty] | 0.42125 | 264.1659 | 0.42125 | 258.9768 | 0.42125 | 267.4844
1.055 | 361.8791 | 0.42254 | 263.9871 | [8 empty] | 0.42254 | 264.7568 | 0.42254 | 259.53 | 0.42254 | 268.1003
1.06 | 361.3298 | 0.42384 | 264.5706 | [8 empty] | 0.42384 | 265.3461 | 0.42384 | 260.0814 | 0.42384 | 268.7145
1.065 | 360.8766 | 0.42514 | 265.1525 | [8 empty] | 0.42514 | 265.9336 | 0.42514 | 260.631 | 0.42514 | 269.3272
1.07 | 360.3273 | 0.42643 | 265.7327 | [8 empty] | 0.42643 | 266.5195 | 0.42643 | 261.1788 | 0.42643 | 269.9383
1.075 | 359.7781 | 0.42773 | 266.3111 | [8 empty] | 0.42773 | 267.1037 | 0.42773 | 261.7248 | 0.42773 | 270.5477
1.08 | 359.2289 | 0.42903 | 266.8879 | [8 empty] | 0.42903 | 267.6862 | 0.42903 | 262.2691 | 0.42903 | 271.1555
1.085 | 358.3916 | 0.43032 | 267.4629 | [8 empty] | 0.43032 | 268.2669 | 0.43032 | 262.8115 | 0.43032 | 271.7617
1.09 | 357.7464 | 0.43162 | 268.0362 | [8 empty] | 0.43162 | 268.846 | 0.43162 | 263.3521 | 0.43162 | 272.3663
1.095 | 357.1011 | 0.43291 | 268.6078 | [8 empty] | 0.43291 | 269.4234 | 0.43291 | 263.891 | 0.43291 | 272.9693
1.1 | 356.5519 | 0.43421 | 269.1777 | [8 empty] | 0.43421 | 269.9991 | 0.43421 | 264.428 | 0.43421 | 273.5706
1.105 | 356.0986 | 0.43551 | 269.7459 | [8 empty] | 0.43551 | 270.5731 | 0.43551 | 264.9633 | 0.43551 | 274.1703
1.11 | 355.3574 | 0.4368 | 270.3124 | [8 empty] | 0.4368 | 271.1454 | 0.4368 | 265.4968 | 0.4368 | 274.7684
1.115 | 354.8081 | 0.4381 | 270.8771 | [8 empty] | 0.4381 | 271.716 | 0.4381 | 266.0285 | 0.4381 | 275.3649
1.12 | 353.9709 | 0.43939 | 271.4401 | [8 empty] | 0.43939 | 272.2849 | 0.43939 | 266.5583 | 0.43939 | 275.9597
1.125 | 353.1336 | 0.44069 | 272.0015 | [8 empty] | 0.44069 | 272.8521 | 0.44069 | 267.0864 | 0.44069 | 276.5529
1.13 | 352.4884 | 0.44199 | 272.5611 | [8 empty] | 0.44199 | 273.4176 | 0.44199 | 267.6127 | 0.44199 | 277.1445
1.135 | 351.6511 | 0.44328 | 273.1189 | [8 empty] | 0.44328 | 273.9814 | 0.44328 | 268.1372 | 0.44328 | 277.7344
1.14 | 350.9099 | 0.44458 | 273.6751 | [8 empty] | 0.44458 | 274.5434 | 0.44458 | 268.6599 | 0.44458 | 278.3227
1.145 | 350.0727 | 0.44587 | 274.2295 | [8 empty] | 0.44587 | 275.1038 | 0.44587 | 269.1808 | 0.44587 | 278.9094
1.15 | 349.4274 | 0.44717 | 274.7822 | [8 empty] | 0.44717 | 275.6625 | 0.44717 | 269.6999 | 0.44717 | 279.4944
1.155 | 348.6862 | 0.44847 | 275.3332 | [8 empty] | 0.44847 | 276.2194 | 0.44847 | 270.2172 | 0.44847 | 280.0778
1.16 | 347.7529 | 0.44976 | 275.8825 | [8 empty] | 0.44976 | 276.7747 | 0.44976 | 270.7327 | 0.44976 | 280.6596
1.165 | 346.9157 | 0.45106 | 276.43 | [8 empty] | 0.45106 | 277.3282 | 0.45106 | 271.2464 | 0.45106 | 281.2397
1.17 | 345.8864 | 0.45236 | 276.9758 | [8 empty] | 0.45236 | 277.88 | 0.45236 | 271.7584 | 0.45236 | 281.8181
1.175 | 344.8572 | 0.45365 | 277.5199 | [8 empty] | 0.45365 | 278.4301 | 0.45365 | 272.2685 | 0.45365 | 282.3949
1.18 | 343.8279 | 0.45495 | 278.0623 | [8 empty] | 0.45495 | 278.9785 | 0.45495 | 272.7768 | 0.45495 | 282.9701
1.185 | 342.6067 | 0.45624 | 278.6029 | [8 empty] | 0.45624 | 279.5252 | 0.45624 | 273.2834 | 0.45624 | 283.5437
1.19 | 341.3854 | 0.45754 | 279.1418 | [8 empty] | 0.45754 | 280.0702 | 0.45754 | 273.7881 | 0.45754 | 284.1155
1.195 | 340.1642 | 0.45884 | 279.679 | [8 empty] | 0.45884 | 280.6134 | 0.45884 | 274.291 | 0.45884 | 284.6858
1.2 | 338.847 | 0.46013 | 280.2144 | [8 empty] | 0.46013 | 281.155 | 0.46013 | 274.7922 | 0.46013 | 285.2543
1.205 | 337.4337 | 0.46143 | 280.7481 | [8 empty] | 0.46143 | 281.6948 | 0.46143 | 275.2915 | 0.46143 | 285.8213
1.21 | 335.7325 | 0.46272 | 281.2801 | [8 empty] | 0.46272 | 282.2329 | 0.46272 | 275.7891 | 0.46272 | 286.3866
1.215 | 333.9352 | 0.46402 | 281.8104 | [8 empty] | 0.46402 | 282.7693 | 0.46402 | 276.2849 | 0.46402 | 286.9502
1.22 | 332.042 | 0.46532 | 282.3389 | [8 empty] | 0.46532 | 283.304 | 0.46532 | 276.7788 | 0.46532 | 287.5122
1.225 | 330.3407 | 0.46661 | 282.8657 | [8 empty] | 0.46661 | 283.8369 | 0.46661 | 277.271 | 0.46661 | 288.0725
1.23 | 328.8315 | 0.46791 | 283.3908 | [8 empty] | 0.46791 | 284.3681 | 0.46791 | 277.7614 | 0.46791 | 288.6311
1.235 | 327.2262 | 0.46921 | 283.9141 | [8 empty] | 0.46921 | 284.8976 | 0.46921 | 278.25 | 0.46921 | 289.1881
1.24 | 325.813 | 0.4705 | 284.4357 | [8 empty] | 0.4705 | 285.4254 | 0.4705 | 278.7368 | 0.4705 | 289.7435
1.245 | 324.3997 | 0.4718 | 284.9556 | [8 empty] | 0.4718 | 285.9515 | 0.4718 | 279.2218 | 0.4718 | 290.2971
1.25 | 322.6985 | 0.47309 | 285.4737 | [8 empty] | 0.47309 | 286.4759 | 0.47309 | 279.705 | 0.47309 | 290.8492
1.255 | 320.6132 | 0.47439 | 285.9901 | [8 empty] | 0.47439 | 286.9985 | 0.47439 | 280.1864 | 0.47439 | 291.3995
1.26 | 318.048 | 0.47569 | 286.5047 | [8 empty] | 0.47569 | 287.5194 | 0.47569 | 280.6661 | 0.47569 | 291.9482
1.265 | 315.5788 | 0.47698 | 287.0177 | [8 empty] | 0.47698 | 288.0386 | 0.47698 | 281.1439 | 0.47698 | 292.4952
1.27 | 313.2055 | 0.47828 | 287.5289 | [8 empty] | 0.47828 | 288.556 | 0.47828 | 281.62 | 0.47828 | 293.0406
1.275 | 310.9283 | 0.47957 | 288.0383 | [8 empty] | 0.47957 | 289.0718 | 0.47957 | 282.0942 | 0.47957 | 293.5843
1.28 | 308.555 | 0.48087 | 288.546 | [8 empty] | 0.48087 | 289.5858 | 0.48087 | 282.5667 | 0.48087 | 294.1263
1.285 | 306.0858 | 0.48217 | 289.052 | [8 empty] | 0.48217 | 290.0981 | 0.48217 | 283.0374 | 0.48217 | 294.6667
1.29 | 303.3285 | 0.48346 | 289.5563 | [8 empty] | 0.48346 | 290.6086 | 0.48346 | 283.5063 | 0.48346 | 295.2054
1.295 | 300.5713 | 0.48476 | 290.0588 | [8 empty] | 0.48476 | 291.1175 | 0.48476 | 283.9734 | 0.48476 | 295.7424
1.3 | 297.142 | 0.48606 | 290.5596 | [8 empty] | 0.48606 | 291.6246 | 0.48606 | 284.4387 | 0.48606 | 296.2778
1.305 | 293.6168 | 0.48735 | 291.0586 | [8 empty] | 0.48735 | 292.1299 | 0.48735 | 284.9022 | 0.48735 | 296.8114
1.31 | 290.0915 | 0.48865 | 291.5559 | [8 empty] | 0.48865 | 292.6336 | 0.48865 | 285.364 | 0.48865 | 297.3434
1.315 | 286.3743 | 0.48994 | 292.0515 | [8 empty] | 0.48994 | 293.1355 | 0.48994 | 285.8239 | 0.48994 | 297.8738
1.32 | 282.465 | 0.49124 | 292.5453 | [8 empty] | 0.49124 | 293.6357 | 0.49124 | 286.2821 | 0.49124 | 298.4024
1.325 | 278.3638 | 0.49254 | 293.0374 | [8 empty] | 0.49254 | 294.1342 | 0.49254 | 286.7385 | 0.49254 | 298.9294
1.33 | 273.4946 | 0.49383 | 293.5277 | [8 empty] | 0.49383 | 294.631 | 0.49383 | 287.1931 | 0.49383 | 299.4547
1.335 | 268.3373 | 0.49513 | 294.0164 | [8 empty] | 0.49513 | 295.126 | 0.49513 | 287.6459 | 0.49513 | 299.9784
1.34 | 262.2201 | 0.49642 | 294.5032 | [8 empty] | 0.49642 | 295.6193 | 0.49642 | 288.097 | 0.49642 | 300.5003
1.345 | 255.6228 | 0.49772 | 294.9884 | [8 empty] | 0.49772 | 296.1108 | 0.49772 | 288.5462 | 0.49772 | 301.0206
1.35 | 247.8736 | 0.49902 | 295.4718 | [8 empty] | 0.49902 | 296.6007 | 0.49902 | 288.9937 | 0.49902 | 301.5392
1.355 | 238.7803 | 0.50031 | 295.9535 | [8 empty] | 0.50031 | 297.0888 | 0.50031 | 289.4394 | 0.50031 | 302.0561
1.36 | 226.4231 | 0.50161 | 296.4334 | [8 empty] | 0.50161 | 297.5752 | 0.50161 | 289.8833 | 0.50161 | 302.5714
1.365 | 212.5298 | 0.50291 | 296.9116 | [8 empty] | 0.50291 | 298.0598 | 0.50291 | 290.3255 | 0.50291 | 303.0849
1.37 | 201.9006 | 0.5042 | 297.388 | [8 empty] | 0.5042 | 298.5428 | 0.5042 | 290.7658 | 0.5042 | 303.5968
1.375 | 195.5913 | 0.5055 | 297.8627 | [8 empty] | 0.5055 | 299.024 | 0.5055 | 291.2044 | 0.5055 | 304.107
1.38 | 194.3701 | 0.50679 | 298.3357 | [8 empty] | 0.50679 | 299.5034 | 0.50679 | 291.6412 | 0.50679 | 304.6155
1.385 | 197.8529 | 0.50809 | 298.807 | [8 empty] | 0.50809 | 299.9812 | 0.50809 | 292.0763 | 0.50809 | 305.1223
1.39 | 200.5676 | 0.50939 | 299.2765 | [8 empty] | 0.50939 | 300.4572 | 0.50939 | 292.5095 | 0.50939 | 305.6275
1.395 | 201.8424 | 0.51068 | 299.7442 | [8 empty] | 0.51068 | 300.9314 | 0.51068 | 292.941 | 0.51068 | 306.131
1.4 | 198.5091 | 0.51198 | 300.2103 | [8 empty] | 0.51198 | 301.404 | 0.51198 | 293.3707 | 0.51198 | 306.6327
1.405 | 193.4479 | 0.51327 | 300.6746 | [8 empty] | 0.51327 | 301.8748 | 0.51327 | 293.7987 | 0.51327 | 307.1328
1.41 | 186.3706 | 0.51457 | 301.1371 | [8 empty] | 0.51457 | 302.3439 | 0.51457 | 294.2248 | 0.51457 | 307.6312
1.415 | 181.4054 | 0.51587 | 301.5979 | [8 empty] | 0.51587 | 302.8113 | 0.51587 | 294.6492 | 0.51587 | 308.1279
1.42 | 179.7041 | 0.51716 | 302.057 | [8 empty] | 0.51716 | 303.2769 | 0.51716 | 295.0719 | 0.51716 | 308.623
1.425 | 182.0349 | 0.51846 | 302.5144 | [8 empty] | 0.51846 | 303.7408 | 0.51846 | 295.4928 | 0.51846 | 309.1163
1.43 | 185.1336 | 0.51976 | 302.97 | [8 empty] | 0.51976 | 304.203 | 0.51976 | 295.9119 | 0.51976 | 309.608
1.435 | 187.1764 | 0.52105 | 303.4239 | [8 empty] | 0.52105 | 304.6635 | 0.52105 | 296.3292 | 0.52105 | 310.0979
1.44 | 187.7791 | 0.52235 | 303.876 | [8 empty] | 0.52235 | 305.1222 | 0.52235 | 296.7448 | 0.52235 | 310.5862
1.445 | 185.9819 | 0.52364 | 304.3264 | [8 empty] | 0.52364 | 305.5792 | 0.52364 | 297.1586 | 0.52364 | 311.0728
1.45 | 183.5127 | 0.52494 | 304.7751 | [8 empty] | 0.52494 | 306.0345 | 0.52494 | 297.5706 | 0.52494 | 311.5577
1.455 | 179.8914 | 0.52624 | 305.222 | [8 empty] | 0.52624 | 306.488 | 0.52624 | 297.9809 | 0.52624 | 312.0409
1.46 | 177.5182 | 0.52753 | 305.6672 | [8 empty] | 0.52753 | 306.9398 | 0.52753 | 298.3894 | 0.52753 | 312.5224
1.465 | 176.1049 | 0.52883 | 306.1107 | [8 empty] | 0.52883 | 307.3899 | 0.52883 | 298.7962 | 0.52883 | 313.0023
1.47 | 176.1317 | 0.53012 | 306.5524 | [8 empty] | 0.53012 | 307.8383 | 0.53012 | 299.2012 | 0.53012 | 313.4804
1.475 | 177.5984 | 0.53142 | 306.9924 | [8 empty] | 0.53142 | 308.2849 | 0.53142 | 299.6045 | 0.53142 | 313.9568
1.48 | 178.6812 | 0.53272 | 307.4307 | [8 empty] | 0.53272 | 308.7298 | 0.53272 | 300.006 | 0.53272 | 314.4316
1.485 | 179.0919 | 0.53401 | 307.8672 | [8 empty] | 0.53401 | 309.173 | 0.53401 | 300.4057 | 0.53401 | 314.9047
1.49 | 178.0627 | 0.53531 | 308.3021 | [8 empty] | 0.53531 | 309.6144 | 0.53531 | 300.8037 | 0.53531 | 315.376
1.495 | 176.0734 | 0.53661 | 308.7351 | [8 empty] | 0.53661 | 310.0542 | 0.53661 | 301.2 | 0.53661 | 315.8457
1.5 | 173.2202 | 0.5379 | 309.1665 | [8 empty] | 0.5379 | 310.4922 | 0.5379 | 301.5944 | 0.5379 | 316.3137
1.505 | 171.711 | 0.5392 | 309.5961 | [8 empty] | 0.5392 | 310.9285 | 0.5392 | 301.9872 | 0.5392 | 316.78
1.51 | 170.6817 | 0.54049 | 310.024 | [8 empty] | 0.54049 | 311.363 | 0.54049 | 302.3782 | 0.54049 | 317.2446
1.515 | 170.4205 | 0.54179 | 310.4501 | [8 empty] | 0.54179 | 311.7958 | 0.54179 | 302.7674 | 0.54179 | 317.7075
1.52 | 170.4472 | 0.54309 | 310.8746 | [8 empty] | 0.54309 | 312.227 | 0.54309 | 303.1549 | 0.54309 | 318.1687
1.525 | 170.57 | 0.54438 | 311.2973 | [8 empty] | 0.54438 | 312.6563 | 0.54438 | 303.5407 | 0.54438 | 318.6282
1.53 | 170.6927 | 0.54568 | 311.7182 | [8 empty] | 0.54568 | 313.084 | 0.54568 | 303.9247 | 0.54568 | 319.086
1.535 | 170.6235 | 0.54697 | 312.1375 | [8 empty] | 0.54697 | 313.5099 | 0.54697 | 304.307 | 0.54697 | 319.5421
1.54 | 169.7862 | 0.54827 | 312.555 | [8 empty] | 0.54827 | 313.9341 | 0.54827 | 304.6875 | 0.54827 | 319.9965
1.545 | 167.509 | 0.54957 | 312.9708 | [8 empty] | 0.54957 | 314.3566 | 0.54957 | 305.0663 | 0.54957 | 320.4493
1.55 | 165.1357 | 0.55086 | 313.3848 | [8 empty] | 0.55086 | 314.7774 | 0.55086 | 305.4434 | 0.55086 | 320.9003
1.555 | 163.4345 | 0.55216 | 313.7972 | [8 empty] | 0.55216 | 315.1965 | 0.55216 | 305.8187 | 0.55216 | 321.3496
1.56 | 163.2692 | 0.55346 | 314.2078 | [8 empty] | 0.55346 | 315.6138 | 0.55346 | 306.1923 | 0.55346 | 321.7973
1.565 | 163.968 | 0.55475 | 314.6167 | [8 empty] | 0.55475 | 316.0294 | 0.55475 | 306.5641 | 0.55475 | 322.2432
1.57 | 164.4748 | 0.55605 | 315.0239 | [8 empty] | 0.55605 | 316.4433 | 0.55605 | 306.9342 | 0.55605 | 322.6875
1.575 | 163.8295 | 0.55734 | 315.4293 | [8 empty] | 0.55734 | 316.8555 | 0.55734 | 307.3026 | 0.55734 | 323.13
1.58 | 162.3203 | 0.55864 | 315.833 | [8 empty] | 0.55864 | 317.2659 | 0.55864 | 307.6693 | 0.55864 | 323.5709
1.585 | 160.523 | 0.55994 | 316.235 | [8 empty] | 0.55994 | 317.6747 | 0.55994 | 308.0342 | 0.55994 | 324.0101
1.59 | 159.3018 | 0.56123 | 316.6353 | [8 empty] | 0.56123 | 318.0817 | 0.56123 | 308.3975 | 0.56123 | 324.4475
1.595 | 158.8485 | 0.56253 | 317.0339 | [8 empty] | 0.56253 | 318.487 | 0.56253 | 308.7589 | 0.56253 | 324.8833
1.6 | 158.3953 | 0.56382 | 317.4307 | [8 empty] | 0.56382 | 318.8906 | 0.56382 | 309.1187 | 0.56382 | 325.3174
1.605 | 157.654 | 0.56512 | 317.8259 | [8 empty] | 0.56512 | 319.2924 | 0.56512 | 309.4767 | 0.56512 | 325.7498
1.61 | 156.4328 | 0.56642 | 318.2193 | [8 empty] | 0.56642 | 319.6926 | 0.56642 | 309.8331 | 0.56642 | 326.1805
1.615 | 155.4035 | 0.56771 | 318.611 | [8 empty] | 0.56771 | 320.091 | 0.56771 | 310.1877 | 0.56771 | 326.6095
1.62 | 155.1423 | 0.56901 | 319.001 | [8 empty] | 0.56901 | 320.4878 | 0.56901 | 310.5406 | 0.56901 | 327.0367
1.625 | 155.169 | 0.57031 | 319.3892 | [8 empty] | 0.57031 | 320.8828 | 0.57031 | 310.8917 | 0.57031 | 327.4623
1.63 | 154.3318 | 0.5716 | 319.7758 | [8 empty] | 0.5716 | 321.2761 | 0.5716 | 311.2412 | 0.5716 | 327.8862
1.635 | 153.2066 | 0.5729 | 320.1606 | [8 empty] | 0.5729 | 321.6677 | 0.5729 | 311.5889 | 0.5729 | 328.3085
1.64 | 151.5053 | 0.57419 | 320.5437 | [8 empty] | 0.57419 | 322.0576 | 0.57419 | 311.935 | 0.57419 | 328.729
1.645 | 149.8041 | 0.57549 | 320.9252 | [8 empty] | 0.57549 | 322.4457 | 0.57549 | 312.2793 | 0.57549 | 329.1478
1.65 | 148.8708 | 0.57679 | 321.3049 | [8 empty] | 0.57679 | 322.8322 | 0.57679 | 312.6219 | 0.57679 | 329.5649
1.655 | 148.7056 | 0.57808 | 321.6829 | [8 empty] | 0.57808 | 323.217 | 0.57808 | 312.9629 | 0.57808 | 329.9803
1.66 | 148.6363 | 0.57938 | 322.0591 | [8 empty] | 0.57938 | 323.6 | 0.57938 | 313.3021 | 0.57938 | 330.3941
1.665 | 147.7031 | 0.58067 | 322.4337 | [8 empty] | 0.58067 | 323.9813 | 0.58067 | 313.6396 | 0.58067 | 330.8061
1.67 | 146.4818 | 0.58197 | 322.8066 | [8 empty] | 0.58197 | 324.361 | 0.58197 | 313.9754 | 0.58197 | 331.2164
1.675 | 144.9726 | 0.58327 | 323.1778 | [8 empty] | 0.58327 | 324.7389 | 0.58327 | 314.3095 | 0.58327 | 331.6251
1.68 | 143.8473 | 0.58456 | 323.5472 | [8 empty] | 0.58456 | 325.1151 | 0.58456 | 314.6419 | 0.58456 | 332.032
1.685 | 143.0101 | 0.58586 | 323.915 | [8 empty] | 0.58586 | 325.4897 | 0.58586 | 314.9727 | 0.58586 | 332.4373
1.69 | 142.4609 | 0.58716 | 324.281 | [8 empty] | 0.58716 | 325.8625 | 0.58716 | 315.3017 | 0.58716 | 332.8409
1.695 | 141.7196 | 0.58845 | 324.6454 | [8 empty] | 0.58845 | 326.2336 | 0.58845 | 315.629 | 0.58845 | 333.2427
1.7 | 140.6904 | 0.58975 | 325.008 | [8 empty] | 0.58975 | 326.603 | 0.58975 | 315.9547 | 0.58975 | 333.6429
1.705 | 139.5651 | 0.59104 | 325.369 | [8 empty] | 0.59104 | 326.9707 | 0.59104 | 316.2786 | 0.59104 | 334.0414
1.71 | 138.5359 | 0.59234 | 325.7282 | [8 empty] | 0.59234 | 327.3367 | 0.59234 | 316.6009 | 0.59234 | 334.4382
1.715 | 137.4106 | 0.59364 | 326.0858 | [8 empty] | 0.59364 | 327.701 | 0.59364 | 316.9215 | 0.59364 | 334.8333
1.72 | 136.5734 | 0.59493 | 326.4416 | [8 empty] | 0.59493 | 328.0637 | 0.59493 | 317.2404 | 0.59493 | 335.2267
1.725 | 135.6401 | 0.59623 | 326.7958 | [8 empty] | 0.59623 | 328.4246 | 0.59623 | 317.5576 | 0.59623 | 335.6184
1.73 | 134.5149 | 0.59752 | 327.1482 | [8 empty] | 0.59752 | 328.7838 | 0.59752 | 317.8732 | 0.59752 | 336.0085
1.735 | 133.3896 | 0.59882 | 327.499 | [8 empty] | 0.59882 | 329.1413 | 0.59882 | 318.1871 | 0.59882 | 336.3968
1.74 | 132.3604 | 0.60012 | 327.8481 | [8 empty] | 0.60012 | 329.4972 | 0.60012 | 318.4993 | 0.60012 | 336.7834
1.745 | 131.2351 | 0.60141 | 328.1955 | [8 empty] | 0.60141 | 329.8513 | 0.60141 | 318.8098 | 0.60141 | 337.1684
1.75 | 130.1099 | 0.60271 | 328.5412 | [8 empty] | 0.60271 | 330.2038 | 0.60271 | 319.1186 | 0.60271 | 337.5516
1.755 | 129.1767 | 0.60401 | 328.8852 | [8 empty] | 0.60401 | 330.5545 | 0.60401 | 319.4258 | 0.60401 | 337.9332
1.76 | 128.4354 | 0.6053 | 329.2275 | [8 empty] | 0.6053 | 330.9036 | 0.6053 | 319.7313 | 0.6053 | 338.3131
1.765 | 127.9822 | 0.6066 | 329.5681 | [8 empty] | 0.6066 | 331.251 | 0.6066 | 320.0352 | 0.6066 | 338.6913
1.77 | 127.1449 | 0.60789 | 329.9071 | [8 empty] | 0.60789 | 331.5967 | 0.60789 | 320.3373 | 0.60789 | 339.0678
1.775 | 125.8277 | 0.60919 | 330.2443 | [8 empty] | 0.60919 | 331.9407 | 0.60919 | 320.6379 | 0.60919 | 339.4426
1.78 | 124.1264 | 0.61049 | 330.5799 | [8 empty] | 0.61049 | 332.283 | 0.61049 | 320.9367 | 0.61049 | 339.8158
1.785 | 122.5212 | 0.61178 | 330.9138 | [8 empty] | 0.61178 | 332.6237 | 0.61178 | 321.2339 | 0.61178 | 340.1872
1.79 | 121.3959 | 0.61308 | 331.246 | [8 empty] | 0.61308 | 332.9626 | 0.61308 | 321.5295 | 0.61308 | 340.557
1.795 | 120.5587 | 0.61437 | 331.5765 | [8 empty] | 0.61437 | 333.2999 | 0.61437 | 321.8234 | 0.61437 | 340.925
1.8 | 119.9134 | 0.61567 | 331.9054 | [8 empty] | 0.61567 | 333.6355 | 0.61567 | 322.1156 | 0.61567 | 341.2914
1.805 | 119.2682 | 0.61697 | 332.2326 | [8 empty] | 0.61697 | 333.9694 | 0.61697 | 322.4062 | 0.61697 | 341.6561
1.81 | 117.951 | 0.61826 | 332.5581 | [8 empty] | 0.61826 | 334.3016 | 0.61826 | 322.6951 | 0.61826 | 342.0191
1.815 | 116.4417 | 0.61956 | 332.8819 | [8 empty] | 0.61956 | 334.6322 | 0.61956 | 322.9824 | 0.61956 | 342.3805
1.82 | 115.1245 | 0.62085 | 333.204 | [8 empty] | 0.62085 | 334.9611 | 0.62085 | 323.2681 | 0.62085 | 342.7401
1.825 | 114.3832 | 0.62215 | 333.5245 | [8 empty] | 0.62215 | 335.2883 | 0.62215 | 323.5521 | 0.62215 | 343.0981
1.83 | 113.546 | 0.62345 | 333.8433 | [8 empty] | 0.62345 | 335.6138 | 0.62345 | 323.8344 | 0.62345 | 343.4544
1.835 | 112.6127 | 0.62474 | 334.1605 | [8 empty] | 0.62474 | 335.9376 | 0.62474 | 324.1152 | 0.62474 | 343.809
1.84 | 111.3915 | 0.62604 | 334.4759 | [8 empty] | 0.62604 | 336.2598 | 0.62604 | 324.3943 | 0.62604 | 344.1619
1.845 | 109.9782 | 0.62734 | 334.7897 | [8 empty] | 0.62734 | 336.5803 | 0.62734 | 324.6717 | 0.62734 | 344.5131
1.85 | 108.757 | 0.62863 | 335.1019 | [8 empty] | 0.62863 | 336.8992 | 0.62863 | 324.9475 | 0.62863 | 344.8627
1.855 | 107.8237 | 0.62993 | 335.4123 | [8 empty] | 0.62993 | 337.2163 | 0.62993 | 325.2217 | 0.62993 | 345.2105
1.86 | 107.1785 | 0.63122 | 335.7211 | [8 empty] | 0.63122 | 337.5318 | 0.63122 | 325.4943 | 0.63122 | 345.5567
1.865 | 106.0532 | 0.63252 | 336.0283 | [8 empty] | 0.63252 | 337.8457 | 0.63252 | 325.7652 | 0.63252 | 345.9013
1.87 | 104.64 | 0.63382 | 336.3338 | [8 empty] | 0.63382 | 338.1578 | 0.63382 | 326.0345 | 0.63382 | 346.2441
1.875 | 103.1308 | 0.63511 | 336.6376 | [8 empty] | 0.63511 | 338.4683 | 0.63511 | 326.3022 | 0.63511 | 346.5853
1.88 | 101.7175 | 0.63641 | 336.9398 | [8 empty] | 0.63641 | 338.7772 | 0.63641 | 326.5683 | 0.63641 | 346.9248
1.885 | 100.8803 | 0.6377 | 337.2403 | [8 empty] | 0.6377 | 339.0844 | 0.6377 | 326.8328 | 0.6377 | 347.2626
1.89 | 100.331 | 0.639 | 337.5391 | [8 empty] | 0.639 | 339.3899 | 0.639 | 327.0956 | 0.639 | 347.5987
1.895 | 99.4938 | 0.6403 | 337.8363 | [8 empty] | 0.6403 | 339.6938 | 0.6403 | 327.3568 | 0.6403 | 347.9332
1.9 | 98.2725 | 0.64159 | 338.1319 | [8 empty] | 0.64159 | 339.996 | 0.64159 | 327.6164 | 0.64159 | 348.266
1.905 | 96.8593 | 0.64289 | 338.4258 | [8 empty] | 0.64289 | 340.2965 | 0.64289 | 327.8744 | 0.64289 | 348.5971
1.91 | 95.446 | 0.64419 | 338.718 | [8 empty] | 0.64419 | 340.5954 | 0.64419 | 328.1308 | 0.64419 | 348.9265
1.915 | 94.2248 | 0.64548 | 339.0086 | [8 empty] | 0.64548 | 340.8927 | 0.64548 | 328.3856 | 0.64548 | 349.2543
1.92 | 93.2915 | 0.64678 | 339.2976 | [8 empty] | 0.64678 | 341.1882 | 0.64678 | 328.6388 | 0.64678 | 349.5804
1.925 | 92.4543 | 0.64807 | 339.5849 | [8 empty] | 0.64807 | 341.4822 | 0.64807 | 328.8904 | 0.64807 | 349.9048
1.93 | 91.2331 | 0.64937 | 339.8706 | [8 empty] | 0.64937 | 341.7745 | 0.64937 | 329.1404 | 0.64937 | 350.2276
1.935 | 90.2038 | 0.65067 | 340.1546 | [8 empty] | 0.65067 | 342.0651 | 0.65067 | 329.3888 | 0.65067 | 350.5487
1.94 | 88.9826 | 0.65196 | 340.437 | [8 empty] | 0.65196 | 342.3541 | 0.65196 | 329.6356 | 0.65196 | 350.8681
1.945 | 87.8573 | 0.65326 | 340.7177 | [8 empty] | 0.65326 | 342.6415 | 0.65326 | 329.8808 | 0.65326 | 351.1859
1.95 | 86.6361 | 0.65455 | 340.9968 | [8 empty] | 0.65455 | 342.9272 | 0.65455 | 330.1244 | 0.65455 | 351.502
1.955 | 85.5108 | 0.65585 | 341.2743 | [8 empty] | 0.65585 | 343.2112 | 0.65585 | 330.3664 | 0.65585 | 351.8164
1.96 | 84.1936 | 0.65715 | 341.5502 | [8 empty] | 0.65715 | 343.4937 | 0.65715 | 330.6068 | 0.65715 | 352.1292
1.965 | 83.1643 | 0.65844 | 341.8244 | [8 empty] | 0.65844 | 343.7744 | 0.65844 | 330.8457 | 0.65844 | 352.4403
1.97 | 82.1351 | 0.65974 | 342.0969 | [8 empty] | 0.65974 | 344.0536 | 0.65974 | 331.083 | 0.65974 | 352.7497
1.975 | 81.1058 | 0.66104 | 342.3679 | [8 empty] | 0.66104 | 344.3311 | 0.66104 | 331.3187 | 0.66104 | 353.0575
1.98 | 79.9806 | 0.66233 | 342.6372 | [8 empty] | 0.66233 | 344.607 | 0.66233 | 331.5528 | 0.66233 | 353.3636
1.985 | 78.8553 | 0.66363 | 342.9049 | [8 empty] | 0.66363 | 344.8812 | 0.66363 | 331.7853 | 0.66363 | 353.6681
1.99 | 77.6341 | 0.66492 | 343.171 | [8 empty] | 0.66492 | 345.1538 | 0.66492 | 332.0163 | 0.66492 | 353.9709
1.995 | 76.4129 | 0.66622 | 343.4354 | [8 empty] | 0.66622 | 345.4248 | 0.66622 | 332.2457 | 0.66622 | 354.272
2 | 74.9996 | 0.66752 | 343.6982 | [8 empty] | 0.66752 | 345.6941 | 0.66752 | 332.4736 | 0.66752 | 354.5715
2.005 | 73.5864 | 0.66881 | 343.9594 | [8 empty] | 0.66881 | 345.9619 | 0.66881 | 332.6998 | 0.66881 | 354.8693
2.01 | 72.0771 | 0.67011 | 344.219 | [8 empty] | 0.67011 | 346.2279 | 0.67011 | 332.9245 | 0.67011 | 355.1655
2.015 | 70.6639 | 0.6714 | 344.477 | [8 empty] | 0.6714 | 346.4924 | 0.6714 | 333.1477 | 0.6714 | 355.46
2.02 | 69.3466 | 0.6727 | 344.7333 | [8 empty] | 0.6727 | 346.7553 | 0.6727 | 333.3693 | 0.6727 | 355.7529
2.025 | 68.3174 | 0.674 | 344.9881 | [8 empty] | 0.674 | 347.0165 | 0.674 | 333.5893 | 0.674 | 356.0441
2.03 | 67.2881 | 0.67529 | 345.2412 | [8 empty] | 0.67529 | 347.2761 | 0.67529 | 333.8078 | 0.67529 | 356.3336
2.035 | 66.0669 | 0.67659 | 345.4927 | [8 empty] | 0.67659 | 347.5341 | 0.67659 | 334.0247 | 0.67659 | 356.6215
2.04 | 64.6536 | 0.67789 | 345.7426 | [8 empty] | 0.67789 | 347.7904 | 0.67789 | 334.2401 | 0.67789 | 356.9078
2.045 | 62.9524 | 0.67918 | 345.9909 | [8 empty] | 0.67918 | 348.0452 | 0.67918 | 334.4539 | 0.67918 | 357.1924
2.05 | 61.2512 | 0.68048 | 346.2376 | [8 empty] | 0.68048 | 348.2983 | 0.68048 | 334.6662 | 0.68048 | 357.4754
2.055 | 60.1259 | 0.68177 | 346.4827 | [8 empty] | 0.68177 | 348.5498 | 0.68177 | 334.8769 | 0.68177 | 357.7567
2.06 | 59.0967 | 0.68307 | 346.7262 | [8 empty] | 0.68307 | 348.7997 | 0.68307 | 335.0861 | 0.68307 | 358.0364
2.065 | 58.1634 | 0.68437 | 346.9681 | [8 empty] | 0.68437 | 349.048 | 0.68437 | 335.2938 | 0.68437 | 358.3144
2.07 | 56.8462 | 0.68566 | 347.2084 | [8 empty] | 0.68566 | 349.2947 | 0.68566 | 335.4999 | 0.68566 | 358.5908
2.075 | 55.2409 | 0.68696 | 347.4471 | [8 empty] | 0.68696 | 349.5398 | 0.68696 | 335.7045 | 0.68696 | 358.8655
2.08 | 53.4437 | 0.68825 | 347.6842 | [8 empty] | 0.68825 | 349.7833 | 0.68825 | 335.9076 | 0.68825 | 359.1386
2.085 | 51.8384 | 0.68955 | 347.9197 | [8 empty] | 0.68955 | 350.0252 | 0.68955 | 336.1091 | 0.68955 | 359.41
2.09 | 50.4252 | 0.69085 | 348.1536 | [8 empty] | 0.69085 | 350.2655 | 0.69085 | 336.3091 | 0.69085 | 359.6799
2.095 | 49.1079 | 0.69214 | 348.386 | [8 empty] | 0.69214 | 350.5041 | 0.69214 | 336.5076 | 0.69214 | 359.948
2.1 | 47.6947 | 0.69344 | 348.6167 | [8 empty] | 0.69344 | 350.7412 | 0.69344 | 336.7046 | 0.69344 | 360.2146
2.105 | 46.2814 | 0.69474 | 348.8459 | [8 empty] | 0.69474 | 350.9767 | 0.69474 | 336.9 | 0.69474 | 360.4795
2.11 | 44.7722 | 0.69603 | 349.0734 | [8 empty] | 0.69603 | 351.2106 | 0.69603 | 337.094 | 0.69603 | 360.7427
2.115 | 43.4549 | 0.69733 | 349.2994 | [8 empty] | 0.69733 | 351.4429 | 0.69733 | 337.2864 | 0.69733 | 361.0044
2.12 | 42.4257 | 0.69862 | 349.5238 | [8 empty] | 0.69862 | 351.6736 | 0.69862 | 337.4773 | 0.69862 | 361.2644
2.125 | 41.2045 | 0.69992 | 349.7466 | [8 empty] | 0.69992 | 351.9027 | 0.69992 | 337.6667 | 0.69992 | 361.5227
2.13 | 39.7912 | 0.70122 | 349.9679 | [8 empty] | 0.70122 | 352.1302 | 0.70122 | 337.8545 | 0.70122 | 361.7795
2.135 | 37.994 | 0.70251 | 350.1876 | [8 empty] | 0.70251 | 352.3562 | 0.70251 | 338.0409 | 0.70251 | 362.0346
2.14 | 36.3887 | 0.70381 | 350.4057 | [8 empty] | 0.70381 | 352.5805 | 0.70381 | 338.2258 | 0.70381 | 362.2881
2.145 | 34.9755 | 0.7051 | 350.6222 | [8 empty] | 0.7051 | 352.8033 | 0.7051 | 338.4092 | 0.7051 | 362.5399
2.15 | 33.6582 | 0.7064 | 350.8372 | [8 empty] | 0.7064 | 353.0245 | 0.7064 | 338.5911 | 0.7064 | 362.7902
2.155 | 32.437 | 0.7077 | 351.0506 | [8 empty] | 0.7077 | 353.2441 | 0.7077 | 338.7715 | 0.7077 | 363.0388
2.16 | 31.0237 | 0.70899 | 351.2624 | [8 empty] | 0.70899 | 353.4621 | 0.70899 | 338.9503 | 0.70899 | 363.2857
2.165 | 29.6105 | 0.71029 | 351.4726 | [8 empty] | 0.71029 | 353.6786 | 0.71029 | 339.1277 | 0.71029 | 363.5311
2.17 | 27.7172 | 0.71159 | 351.6813 | [8 empty] | 0.71159 | 353.8935 | 0.71159 | 339.3037 | 0.71159 | 363.7748
2.175 | 26.016 | 0.71288 | 351.8885 | [8 empty] | 0.71288 | 354.1068 | 0.71288 | 339.4781 | 0.71288 | 364.0169
2.18 | 24.5068 | 0.71418 | 352.094 | [8 empty] | 0.71418 | 354.3185 | 0.71418 | 339.651 | 0.71418 | 364.2574
2.185 | 22.9975 | 0.71547 | 352.2981 | [8 empty] | 0.71547 | 354.5287 | 0.71547 | 339.8225 | 0.71547 | 364.4963
2.19 | 21.8723 | 0.71677 | 352.5005 | [8 empty] | 0.71677 | 354.7373 | 0.71677 | 339.9925 | 0.71677 | 364.7336
2.195 | 20.555 | 0.71807 | 352.7014 | [8 empty] | 0.71807 | 354.9443 | 0.71807 | 340.161 | 0.71807 | 364.9692
2.2 | 19.2378 | 0.71936 | 352.9008 | [8 empty] | 0.71936 | 355.1498 | 0.71936 | 340.328 | 0.71936 | 365.2033
2.205 | 17.9205 | 0.72066 | 353.0986 | [8 empty] | 0.72066 | 355.3537 | 0.72066 | 340.4936 | 0.72066 | 365.4357
2.21 | 16.5073 | 0.72195 | 353.2949 | [8 empty] | 0.72195 | 355.556 | 0.72195 | 340.6577 | 0.72195 | 365.6665
2.215 | 14.998 | 0.72325 | 353.4896 | [8 empty] | 0.72325 | 355.7568 | 0.72325 | 340.8203 | 0.72325 | 365.8957
2.22 | 13.4888 | 0.72455 | 353.6828 | [8 empty] | 0.72455 | 355.9561 | 0.72455 | 340.9815 | 0.72455 | 366.1233
2.225 | 11.9795 | 0.72584 | 353.8744 | [8 empty] | 0.72584 | 356.1537 | 0.72584 | 341.1412 | 0.72584 | 366.3492
2.23 | 10.6623 | 0.72714 | 354.0645 | [8 empty] | 0.72714 | 356.3499 | 0.72714 | 341.2995 | 0.72714 | 366.5736
2.235 | 9.3451 | 0.72844 | 354.2531 | [8 empty] | 0.72844 | 356.5444 | 0.72844 | 341.4563 | 0.72844 | 366.7964
2.24 | 7.7398 | 0.72973 | 354.4401 | [8 empty] | 0.72973 | 356.7375 | 0.72973 | 341.6116 | 0.72973 | 367.0176
2.245 | 6.2306 | 0.73103 | 354.6256 | [8 empty] | 0.73103 | 356.9289 | 0.73103 | 341.7655 | 0.73103 | 367.2371
2.25 | 4.8173 | 0.73232 | 354.8095 | [8 empty] | 0.73232 | 357.1189 | 0.73232 | 341.9179 | 0.73232 | 367.4551
2.255 | 3.5001 | 0.73362 | 354.9919 | [8 empty] | 0.73362 | 357.3073 | 0.73362 | 342.0689 | 0.73362 | 367.6715
2.26 | 1.9908 | 0.73492 | 355.1728 | [8 empty] | 0.73492 | 357.4941 | 0.73492 | 342.2185 | 0.73492 | 367.8862
2.265 | 0.48157 | 0.73621 | 355.3522 | [8 empty] | 0.73621 | 357.6794 | 0.73621 | 342.3666 | 0.73621 | 368.0994
2.27 | -1.1237 | 0.73751 | 355.53 | [8 empty] | 0.73751 | 357.8632 | 0.73751 | 342.5133 | 0.73751 | 368.311
2.275 | -2.7289 | 0.7388 | 355.7064 | [8 empty] | 0.7388 | 358.0454 | 0.7388 | 342.6585 | 0.7388 | 368.5209
2.28 | -4.2382 | 0.7401 | 355.8812 | [8 empty] | 0.7401 | 358.2261 | 0.7401 | 342.8023 | 0.7401 | 368.7293
2.285 | -5.5554 | 0.7414 | 356.0544 | [8 empty] | 0.7414 | 358.4053 | 0.7414 | 342.9447 | 0.7414 | 368.9361
2.29 | -6.8727 | 0.74269 | 356.2262 | [8 empty] | 0.74269 | 358.5829 | 0.74269 | 343.0856 | 0.74269 | 369.1413
2.295 | -8.1899 | 0.74399 | 356.3964 | [8 empty] | 0.74399 | 358.759 | 0.74399 | 343.2251 | 0.74399 | 369.3449
2.3 | -9.6031 | 0.74529 | 356.5652 | [8 empty] | 0.74529 | 358.9336 | 0.74529 | 343.3632 | 0.74529 | 369.547
2.305 | -11.3044 | 0.74658 | 356.7324 | [8 empty] | 0.74658 | 359.1066 | 0.74658 | 343.4999 | 0.74658 | 369.7474
2.31 | -12.9096 | 0.74788 | 356.8981 | [8 empty] | 0.74788 | 359.2781 | 0.74788 | 343.6351 | 0.74788 | 369.9463
2.315 | -14.3229 | 0.74917 | 357.0623 | [8 empty] | 0.74917 | 359.4481 | 0.74917 | 343.7689 | 0.74917 | 370.1435
2.32 | -15.5441 | 0.75047 | 357.225 | [8 empty] | 0.75047 | 359.6166 | 0.75047 | 343.9014 | 0.75047 | 370.3392
2.325 | -16.8614 | 0.75177 | 357.3863 | [8 empty] | 0.75177 | 359.7836 | 0.75177 | 344.0324 | 0.75177 | 370.5333
2.33 | -18.2746 | 0.75306 | 357.546 | [8 empty] | 0.75306 | 359.949 | 0.75306 | 344.162 | 0.75306 | 370.7259
2.335 | -19.7839 | 0.75436 | 357.7042 | [8 empty] | 0.75436 | 360.113 | 0.75436 | 344.2902 | 0.75436 | 370.9168
2.34 | -21.3891 | 0.75565 | 357.8609 | [8 empty] | 0.75565 | 360.2754 | 0.75565 | 344.4169 | 0.75565 | 371.1062
2.345 | -22.8984 | 0.75695 | 358.0161 | [8 empty] | 0.75695 | 360.4363 | 0.75695 | 344.5423 | 0.75695 | 371.294
2.35 | -24.1196 | 0.75825 | 358.1698 | [8 empty] | 0.75825 | 360.5957 | 0.75825 | 344.6663 | 0.75825 | 371.4802
2.355 | -25.3408 | 0.75954 | 358.322 | [8 empty] | 0.75954 | 360.7536 | 0.75954 | 344.7889 | 0.75954 | 371.6649
2.36 | -26.8501 | 0.76084 | 358.4728 | [8 empty] | 0.76084 | 360.91 | 0.76084 | 344.9101 | 0.76084 | 371.848
2.365 | -28.3593 | 0.76214 | 358.622 | [8 empty] | 0.76214 | 361.0649 | 0.76214 | 345.0299 | 0.76214 | 372.0295
2.37 | -29.8686 | 0.76343 | 358.7698 | [8 empty] | 0.76343 | 361.2183 | 0.76343 | 345.1483 | 0.76343 | 372.2094
2.375 | -31.3778 | 0.76473 | 358.9161 | [8 empty] | 0.76473 | 361.3702 | 0.76473 | 345.2654 | 0.76473 | 372.3878
2.38 | -32.6951 | 0.76602 | 359.0609 | [8 empty] | 0.76602 | 361.5206 | 0.76602 | 345.381 | 0.76602 | 372.5646
2.385 | -34.0123 | 0.76732 | 359.2042 | [8 empty] | 0.76732 | 361.6695 | 0.76732 | 345.4953 | 0.76732 | 372.7399
2.39 | -35.5216 | 0.76862 | 359.3461 | [8 empty] | 0.76862 | 361.8169 | 0.76862 | 345.6082 | 0.76862 | 372.9136
2.395 | -37.0308 | 0.76991 | 359.4865 | [8 empty] | 0.76991 | 361.9628 | 0.76991 | 345.7197 | 0.76991 | 373.0857
2.4 | -38.4441 | 0.77121 | 359.6254 | [8 empty] | 0.77121 | 362.1073 | 0.77121 | 345.8298 | 0.77121 | 373.2563
2.405 | -39.7613 | 0.7725 | 359.7628 | [8 empty] | 0.7725 | 362.2502 | 0.7725 | 345.9386 | 0.7725 | 373.4253
2.41 | -41.0786 | 0.7738 | 359.8988 | [8 empty] | 0.7738 | 362.3917 | 0.7738 | 346.046 | 0.7738 | 373.5928
2.415 | -42.3958 | 0.7751 | 360.0333 | [8 empty] | 0.7751 | 362.5317 | 0.7751 | 346.152 | 0.7751 | 373.7587
2.42 | -43.7131 | 0.77639 | 360.1664 | [8 empty] | 0.77639 | 362.6702 | 0.77639 | 346.2567 | 0.77639 | 373.9231
2.425 | -45.1263 | 0.77769 | 360.298 | [8 empty] | 0.77769 | 362.8072 | 0.77769 | 346.36 | 0.77769 | 374.0859
2.43 | -46.4435 | 0.77899 | 360.4281 | [8 empty] | 0.77899 | 362.9427 | 0.77899 | 346.462 | 0.77899 | 374.2472
2.435 | -48.0488 | 0.78028 | 360.5567 | [8 empty] | 0.78028 | 363.0768 | 0.78028 | 346.5626 | 0.78028 | 374.4069
2.44 | -49.366 | 0.78158 | 360.684 | [8 empty] | 0.78158 | 363.2094 | 0.78158 | 346.6618 | 0.78158 | 374.565
2.445 | -50.5873 | 0.78287 | 360.8097 | [8 empty] | 0.78287 | 363.3405 | 0.78287 | 346.7597 | 0.78287 | 374.7217
2.45 | -52.0005 | 0.78417 | 360.934 | [8 empty] | 0.78417 | 363.4702 | 0.78417 | 346.8562 | 0.78417 | 374.8768
2.455 | -53.4138 | 0.78547 | 361.0569 | [8 empty] | 0.78547 | 363.5984 | 0.78547 | 346.9514 | 0.78547 | 375.0303
2.46 | -54.443 | 0.78676 | 361.1783 | [8 empty] | 0.78676 | 363.7251 | 0.78676 | 347.0453 | 0.78676 | 375.1823
2.465 | -55.7603 | 0.78806 | 361.2983 | [8 empty] | 0.78806 | 363.8504 | 0.78806 | 347.1378 | 0.78806 | 375.3328
2.47 | -57.0775 | 0.78935 | 361.4168 | [8 empty] | 0.78935 | 363.9742 | 0.78935 | 347.229 | 0.78935 | 375.4817
2.475 | -58.3948 | 0.79065 | 361.5339 | [8 empty] | 0.79065 | 364.0965 | 0.79065 | 347.3188 | 0.79065 | 375.6291
2.48 | -59.808 | 0.79195 | 361.6495 | [8 empty] | 0.79195 | 364.2174 | 0.79195 | 347.4073 | 0.79195 | 375.7749
2.485 | -61.2212 | 0.79324 | 361.7637 | [8 empty] | 0.79324 | 364.3369 | 0.79324 | 347.4945 | 0.79324 | 375.9193
2.49 | -62.5385 | 0.79454 | 361.8765 | [8 empty] | 0.79454 | 364.4548 | 0.79454 | 347.5803 | 0.79454 | 376.0621
2.495 | -63.7597 | 0.79583 | 361.9879 | [8 empty] | 0.79583 | 364.5714 | 0.79583 | 347.6649 | 0.79583 | 376.2033
2.5 | -64.885 | 0.79713 | 362.0978 | [8 empty] | 0.79713 | 364.6865 | 0.79713 | 347.7481 | 0.79713 | 376.3431
2.505 | -66.2022 | 0.79843 | 362.2063 | [8 empty] | 0.79843 | 364.8001 | 0.79843 | 347.83 | 0.79843 | 376.4813
2.51 | -67.3275 | 0.79972 | 362.3133 | [8 empty] | 0.79972 | 364.9123 | 0.79972 | 347.9105 | 0.79972 | 376.618
2.515 | -68.7407 | 0.80102 | 362.419 | [8 empty] | 0.80102 | 365.0231 | 0.80102 | 347.9898 | 0.80102 | 376.7531
2.52 | -70.058 | 0.80232 | 362.5232 | [8 empty] | 0.80232 | 365.1324 | 0.80232 | 348.0677 | 0.80232 | 376.8868
2.525 | -71.2792 | 0.80361 | 362.626 | [8 empty] | 0.80361 | 365.2403 | 0.80361 | 348.1443 | 0.80361 | 377.0189
2.53 | -72.6925 | 0.80491 | 362.7274 | [8 empty] | 0.80491 | 365.3467 | 0.80491 | 348.2196 | 0.80491 | 377.1495
2.535 | -74.0097 | 0.8062 | 362.8274 | [8 empty] | 0.8062 | 365.4517 | 0.8062 | 348.2937 | 0.8062 | 377.2786
2.54 | -75.3269 | 0.8075 | 362.9259 | [8 empty] | 0.8075 | 365.5553 | 0.8075 | 348.3664 | 0.8075 | 377.4062
2.545 | -76.5482 | 0.8088 | 363.0231 | [8 empty] | 0.8088 | 365.6574 | 0.8088 | 348.4378 | 0.8088 | 377.5323
2.55 | -77.6734 | 0.81009 | 363.1188 | [8 empty] | 0.81009 | 365.7582 | 0.81009 | 348.5079 | 0.81009 | 377.6569
2.555 | -78.7987 | 0.81139 | 363.2131 | [8 empty] | 0.81139 | 365.8575 | 0.81139 | 348.5767 | 0.81139 | 377.7799
2.56 | -79.8279 | 0.81268 | 363.3061 | [8 empty] | 0.81268 | 365.9553 | 0.81268 | 348.6442 | 0.81268 | 377.9015
2.565 | -81.0492 | 0.81398 | 363.3976 | [8 empty] | 0.81398 | 366.0518 | 0.81398 | 348.7105 | 0.81398 | 378.0215
2.57 | -82.3664 | 0.81528 | 363.4877 | [8 empty] | 0.81528 | 366.1468 | 0.81528 | 348.7754 | 0.81528 | 378.1401
2.575 | -83.5877 | 0.81657 | 363.5765 | [8 empty] | 0.81657 | 366.2404 | 0.81657 | 348.8391 | 0.81657 | 378.2571
2.58 | -84.7129 | 0.81787 | 363.6638 | [8 empty] | 0.81787 | 366.3326 | 0.81787 | 348.9015 | 0.81787 | 378.3727
2.585 | -85.8382 | 0.81917 | 363.7497 | [8 empty] | 0.81917 | 366.4234 | 0.81917 | 348.9626 | 0.81917 | 378.4867
2.59 | -86.8674 | 0.82046 | 363.8343 | [8 empty] | 0.82046 | 366.5128 | 0.82046 | 349.0224 | 0.82046 | 378.5992
2.595 | -87.8967 | 0.82176 | 363.9175 | [8 empty] | 0.82176 | 366.6008 | 0.82176 | 349.081 | 0.82176 | 378.7103
2.6 | -88.9259 | 0.82305 | 363.9992 | [8 empty] | 0.82305 | 366.6874 | 0.82305 | 349.1383 | 0.82305 | 378.8198
2.605 | -90.0512 | 0.82435 | 364.0796 | [8 empty] | 0.82435 | 366.7725 | 0.82435 | 349.1943 | 0.82435 | 378.9279
2.61 | -91.1764 | 0.82565 | 364.1587 | [8 empty] | 0.82565 | 366.8563 | 0.82565 | 349.249 | 0.82565 | 379.0345
2.615 | -92.3016 | 0.82694 | 364.2363 | [8 empty] | 0.82694 | 366.9386 | 0.82694 | 349.3025 | 0.82694 | 379.1396
2.62 | -93.2349 | 0.82824 | 364.3126 | [8 empty] | 0.82824 | 367.0196 | 0.82824 | 349.3547 | 0.82824 | 379.2432
2.625 | -94.0721 | 0.82953 | 364.3874 | [8 empty] | 0.82953 | 367.0992 | 0.82953 | 349.4057 | 0.82953 | 379.3453
2.63 | -95.1014 | 0.83083 | 364.4609 | [8 empty] | 0.83083 | 367.1774 | 0.83083 | 349.4554 | 0.83083 | 379.4459
2.635 | -96.1306 | 0.83213 | 364.5331 | [8 empty] | 0.83213 | 367.2541 | 0.83213 | 349.5038 | 0.83213 | 379.5451
2.64 | -97.0639 | 0.83342 | 364.6039 | [8 empty] | 0.83342 | 367.3295 | 0.83342 | 349.551 | 0.83342 | 379.6428
2.645 | -97.9011 | 0.83472 | 364.6733 | [8 empty] | 0.83472 | 367.4035 | 0.83472 | 349.597 | 0.83472 | 379.739
2.65 | -98.9304 | 0.83602 | 364.7413 | [8 empty] | 0.83602 | 367.4762 | 0.83602 | 349.6417 | 0.83602 | 379.8337
2.655 | -99.8636 | 0.83731 | 364.808 | [8 empty] | 0.83731 | 367.5474 | 0.83731 | 349.6852 | 0.83731 | 379.9269
2.66 | -100.7969 | 0.83861 | 364.8733 | [8 empty] | 0.83861 | 367.6173 | 0.83861 | 349.7274 | 0.83861 | 380.0187
2.665 | -101.7301 | 0.8399 | 364.9373 | [8 empty] | 0.8399 | 367.6857 | 0.8399 | 349.7684 | 0.8399 | 380.109
2.67 | -102.6633 | 0.8412 | 364.9999 | [8 empty] | 0.8412 | 367.7528 | 0.8412 | 349.8081 | 0.8412 | 380.1979
2.675 | -103.4046 | 0.8425 | 365.0611 | [8 empty] | 0.8425 | 367.8186 | 0.8425 | 349.8466 | 0.8425 | 380.2852
2.68 | -104.2418 | 0.84379 | 365.121 | [8 empty] | 0.84379 | 367.8829 | 0.84379 | 349.8839 | 0.84379 | 380.3711
2.685 | -104.9831 | 0.84509 | 365.1796 | [8 empty] | 0.84509 | 367.9459 | 0.84509 | 349.9199 | 0.84509 | 380.4556
2.69 | -105.8203 | 0.84638 | 365.2368 | [8 empty] | 0.84638 | 368.0076 | 0.84638 | 349.9548 | 0.84638 | 380.5386
2.695 | -106.5616 | 0.84768 | 365.2927 | [8 empty] | 0.84768 | 368.0678 | 0.84768 | 349.9884 | 0.84768 | 380.6201
2.7 | -107.3988 | 0.84898 | 365.3472 | [8 empty] | 0.84898 | 368.1267 | 0.84898 | 350.0208 | 0.84898 | 380.7001
2.705 | -108.2361 | 0.85027 | 365.4004 | [8 empty] | 0.85027 | 368.1842 | 0.85027 | 350.0519 | 0.85027 | 380.7788
2.71 | -109.0733 | 0.85157 | 365.4522 | [8 empty] | 0.85157 | 368.2404 | 0.85157 | 350.0819 | 0.85157 | 380.8559
2.715 | -109.9106 | 0.85287 | 365.5027 | [8 empty] | 0.85287 | 368.2952 | 0.85287 | 350.1106 | 0.85287 | 380.9316
2.72 | -110.7478 | 0.85416 | 365.5519 | [8 empty] | 0.85416 | 368.3487 | 0.85416 | 350.1381 | 0.85416 | 381.0059
2.725 | -111.489 | 0.85546 | 365.5998 | [8 empty] | 0.85546 | 368.4008 | 0.85546 | 350.1644 | 0.85546 | 381.0787
2.73 | -112.2303 | 0.85675 | 365.6463 | [8 empty] | 0.85675 | 368.4516 | 0.85675 | 350.1895 | 0.85675 | 381.15
2.735 | -112.8755 | 0.85805 | 365.6915 | [8 empty] | 0.85805 | 368.501 | 0.85805 | 350.2134 | 0.85805 | 381.2199
2.74 | -113.6168 | 0.85935 | 365.7354 | [8 empty] | 0.85935 | 368.549 | 0.85935 | 350.2362 | 0.85935 | 381.2884
2.745 | -114.358 | 0.86064 | 365.7779 | [8 empty] | 0.86064 | 368.5958 | 0.86064 | 350.2577 | 0.86064 | 381.3554
2.75 | -115.0033 | 0.86194 | 365.8191 | [8 empty] | 0.86194 | 368.6411 | 0.86194 | 350.278 | 0.86194 | 381.421
2.755 | -115.8405 | 0.86323 | 365.859 | [8 empty] | 0.86323 | 368.6852 | 0.86323 | 350.2971 | 0.86323 | 381.4852
2.76 | -116.4858 | 0.86453 | 365.8976 | [8 empty] | 0.86453 | 368.7279 | 0.86453 | 350.315 | 0.86453 | 381.5479
2.765 | -116.939 | 0.86583 | 365.9349 | [8 empty] | 0.86583 | 368.7693 | 0.86583 | 350.3318 | 0.86583 | 381.6092
2.77 | -117.5843 | 0.86712 | 365.9709 | [8 empty] | 0.86712 | 368.8093 | 0.86712 | 350.3473 | 0.86712 | 381.669
2.775 | -118.2295 | 0.86842 | 366.0056 | [8 empty] | 0.86842 | 368.848 | 0.86842 | 350.3617 | 0.86842 | 381.7274
2.78 | -118.9708 | 0.86972 | 366.0389 | [8 empty] | 0.86972 | 368.8854 | 0.86972 | 350.3749 | 0.86972 | 381.7844
2.785 | -119.616 | 0.87101 | 366.071 | [8 empty] | 0.87101 | 368.9214 | 0.87101 | 350.3869 | 0.87101 | 381.84
2.79 | -120.1652 | 0.87231 | 366.1017 | [8 empty] | 0.87231 | 368.9562 | 0.87231 | 350.3977 | 0.87231 | 381.8941
2.795 | -120.7145 | 0.8736 | 366.1312 | [8 empty] | 0.8736 | 368.9896 | 0.8736 | 350.4074 | 0.8736 | 381.9468
2.8 | -121.2637 | 0.8749 | 366.1594 | [8 empty] | 0.8749 | 369.0217 | 0.8749 | 350.4159 | 0.8749 | 381.9981
2.805 | -121.813 | 0.8762 | 366.1862 | [8 empty] | 0.8762 | 369.0524 | 0.8762 | 350.4232 | 0.8762 | 382.048
2.81 | -122.3622 | 0.87749 | 366.2118 | [8 empty] | 0.87749 | 369.0819 | 0.87749 | 350.4294 | 0.87749 | 382.0965
2.815 | -123.0075 | 0.87879 | 366.2361 | [8 empty] | 0.87879 | 369.11 | 0.87879 | 350.4344 | 0.87879 | 382.1435
2.82 | -123.4607 | 0.88008 | 366.2591 | [8 empty] | 0.88008 | 369.1369 | 0.88008 | 350.4382 | 0.88008 | 382.1892
2.825 | -124.01 | 0.88138 | 366.2808 | [8 empty] | 0.88138 | 369.1624 | 0.88138 | 350.4409 | 0.88138 | 382.2334
2.83 | -124.4632 | 0.88268 | 366.3012 | [8 empty] | 0.88268 | 369.1866 | 0.88268 | 350.4424 | 0.88268 | 382.2762
2.835 | -125.0125 | 0.88397 | 366.3204 | [8 empty] | 0.88397 | 369.2095 | 0.88397 | 350.4428 | 0.88397 | 382.3176
2.84 | -125.4657 | 0.88527 | 366.3382 | [8 empty] | 0.88527 | 369.2311 | 0.88527 | 350.442 | 0.88527 | 382.3576
2.845 | -125.823 | 0.88657 | 366.3548 | [8 empty] | 0.88657 | 369.2514 | 0.88657 | 350.44 | 0.88657 | 382.3962
2.85 | -126.2762 | 0.88786 | 366.3701 | [8 empty] | 0.88786 | 369.2704 | 0.88786 | 350.437 | 0.88786 | 382.4334
2.855 | -126.8254 | 0.88916 | 366.3842 | [8 empty] | 0.88916 | 369.2882 | 0.88916 | 350.4328 | 0.88916 | 382.4692
2.86 | -127.2787 | 0.89045 | 366.397 | [8 empty] | 0.89045 | 369.3046 | 0.89045 | 350.4274 | 0.89045 | 382.5036
2.865 | -127.7319 | 0.89175 | 366.4085 | [8 empty] | 0.89175 | 369.3197 | 0.89175 | 350.4209 | 0.89175 | 382.5366
2.87 | -128.0892 | 0.89305 | 366.4187 | [8 empty] | 0.89305 | 369.3336 | 0.89305 | 350.4133 | 0.89305 | 382.5682
2.875 | -128.5424 | 0.89434 | 366.4277 | [8 empty] | 0.89434 | 369.3461 | 0.89434 | 350.4045 | 0.89434 | 382.5984
2.88 | -128.8997 | 0.89564 | 366.4354 | [8 empty] | 0.89564 | 369.3574 | 0.89564 | 350.3946 | 0.89564 | 382.6272
2.885 | -129.2569 | 0.89693 | 366.4419 | [8 empty] | 0.89693 | 369.3674 | 0.89693 | 350.3836 | 0.89693 | 382.6546
2.89 | -129.6142 | 0.89823 | 366.4471 | [8 empty] | 0.89823 | 369.3761 | 0.89823 | 350.3714 | 0.89823 | 382.6807
2.895 | -129.9714 | 0.89953 | 366.451 | [8 empty] | 0.89953 | 369.3835 | 0.89953 | 350.3582 | 0.89953 | 382.7054
2.9 | -130.2327 | 0.90082 | 366.4537 | [8 empty] | 0.90082 | 369.3897 | 0.90082 | 350.3438 | 0.90082 | 382.7286
2.905 | -130.5899 | 0.90212 | 366.4552 | [8 empty] | 0.90212 | 369.3946 | 0.90212 | 350.3283 | 0.90212 | 382.7505
2.91 | -130.9471 | 0.90342 | 366.4554 | [8 empty] | 0.90342 | 369.3982 | 0.90342 | 350.3116 | 0.90342 | 382.7711
2.915 | -131.2084 | 0.90471 | 366.4543 | [8 empty] | 0.90471 | 369.4005 | 0.90471 | 350.2939 | 0.90471 | 382.7902
2.92 | -131.4696 | 0.90601 | 366.452 | [8 empty] | 0.90601 | 369.4016 | 0.90601 | 350.275 | 0.90601 | 382.808
2.925 | -131.7309 | 0.9073 | 366.4485 | [8 empty] | 0.9073 | 369.4014 | 0.9073 | 350.2551 | 0.9073 | 382.8244
2.93 | -132.0881 | 0.9086 | 366.4437 | [8 empty] | 0.9086 | 369.4 | 0.9086 | 350.234 | 0.9086 | 382.8394
2.935 | -132.2534 | 0.9099 | 366.4377 | [8 empty] | 0.9099 | 369.3973 | 0.9099 | 350.2118 | 0.9099 | 382.853
2.94 | -132.4186 | 0.91119 | 366.4305 | [8 empty] | 0.91119 | 369.3933 | 0.91119 | 350.1886 | 0.91119 | 382.8653
2.945 | -132.7759 | 0.91249 | 366.422 | [8 empty] | 0.91249 | 369.3881 | 0.91249 | 350.1642 | 0.91249 | 382.8763
2.95 | -133.1331 | 0.91378 | 366.4123 | [8 empty] | 0.91378 | 369.3816 | 0.91378 | 350.1387 | 0.91378 | 382.8858
2.955 | -133.2024 | 0.91508 | 366.4014 | [8 empty] | 0.91508 | 369.3739 | 0.91508 | 350.1122 | 0.91508 | 382.894
2.96 | -133.3676 | 0.91638 | 366.3893 | [8 empty] | 0.91638 | 369.3649 | 0.91638 | 350.0845 | 0.91638 | 382.9009
2.965 | -133.6289 | 0.91767 | 366.3759 | [8 empty] | 0.91767 | 369.3547 | 0.91767 | 350.0558 | 0.91767 | 382.9063
2.97 | -133.8901 | 0.91897 | 366.3613 | [8 empty] | 0.91897 | 369.3432 | 0.91897 | 350.0259 | 0.91897 | 382.9105
2.975 | -134.1513 | 0.92027 | 366.3455 | [8 empty] | 0.92027 | 369.3305 | 0.92027 | 349.995 | 0.92027 | 382.9133
2.98 | -134.1246 | 0.92156 | 366.3285 | [8 empty] | 0.92156 | 369.3166 | 0.92156 | 349.963 | 0.92156 | 382.9147
2.985 | -134.3858 | 0.92286 | 366.3103 | [8 empty] | 0.92286 | 369.3014 | 0.92286 | 349.93 | 0.92286 | 382.9147
2.99 | -134.6471 | 0.92415 | 366.2908 | [8 empty] | 0.92415 | 369.285 | 0.92415 | 349.8958 | 0.92415 | 382.9135
2.995 | -134.8123 | 0.92545 | 366.2702 | [8 empty] | 0.92545 | 369.2673 | 0.92545 | 349.8606 | 0.92545 | 382.9109
3 | -135.0736 | 0.92675 | 366.2483 | [8 empty] | 0.92675 | 369.2485 | 0.92675 | 349.8243 | 0.92675 | 382.9069
3.005 | -135.3348 | 0.92804 | 366.2253 | [8 empty] | 0.92804 | 369.2283 | 0.92804 | 349.7869 | 0.92804 | 382.9016
3.01 | -135.5961 | 0.92934 | 366.201 | [8 empty] | 0.92934 | 369.207 | 0.92934 | 349.7485 | 0.92934 | 382.8949
3.015 | -135.7613 | 0.93063 | 366.1755 | [8 empty] | 0.93063 | 369.1844 | 0.93063 | 349.709 | 0.93063 | 382.887
3.02 | -136.1186 | 0.93193 | 366.1489 | [8 empty] | 0.93193 | 369.1607 | 0.93193 | 349.6685 | 0.93193 | 382.8777
3.025 | -136.3798 | 0.93323 | 366.121 | [8 empty] | 0.93323 | 369.1356 | 0.93323 | 349.6268 | 0.93323 | 382.867
3.03 | -136.353 | 0.93452 | 366.092 | [8 empty] | 0.93452 | 369.1094 | 0.93452 | 349.5842 | 0.93452 | 382.855
3.035 | -136.4223 | 0.93582 | 366.0617 | [8 empty] | 0.93582 | 369.082 | 0.93582 | 349.5404 | 0.93582 | 382.8417
3.04 | -136.3955 | 0.93712 | 366.0303 | [8 empty] | 0.93712 | 369.0533 | 0.93712 | 349.4957 | 0.93712 | 382.8271
3.045 | -136.3688 | 0.93841 | 365.9977 | [8 empty] | 0.93841 | 369.0235 | 0.93841 | 349.4498 | 0.93841 | 382.8111
3.05 | -136.342 | 0.93971 | 365.9639 | [8 empty] | 0.93971 | 368.9924 | 0.93971 | 349.403 | 0.93971 | 382.7938
3.055 | -136.4113 | 0.941 | 365.9289 | [8 empty] | 0.941 | 368.9601 | 0.941 | 349.355 | 0.941 | 382.7752
3.06 | -136.5765 | 0.9423 | 365.8927 | [8 empty] | 0.9423 | 368.9266 | 0.9423 | 349.3061 | 0.9423 | 382.7553
3.065 | -136.8378 | 0.9436 | 365.8554 | [8 empty] | 0.9436 | 368.892 | 0.9436 | 349.2561 | 0.9436 | 382.734
3.07 | -137.003 | 0.94489 | 365.8169 | [8 empty] | 0.94489 | 368.8561 | 0.94489 | 349.205 | 0.94489 | 382.7115
3.075 | -136.8803 | 0.94619 | 365.7772 | [8 empty] | 0.94619 | 368.819 | 0.94619 | 349.1529 | 0.94619 | 382.6876
3.08 | -136.9495 | 0.94748 | 365.7364 | [8 empty] | 0.94748 | 368.7807 | 0.94748 | 349.0998 | 0.94748 | 382.6624
3.085 | -136.9228 | 0.94878 | 365.6943 | [8 empty] | 0.94878 | 368.7412 | 0.94878 | 349.0457 | 0.94878 | 382.6359
3.09 | -136.992 | 0.95008 | 365.6512 | [8 empty] | 0.95008 | 368.7006 | 0.95008 | 348.9905 | 0.95008 | 382.6081
3.095 | -136.8692 | 0.95137 | 365.6068 | [8 empty] | 0.95137 | 368.6587 | 0.95137 | 348.9343 | 0.95137 | 382.579
3.1 | -136.9385 | 0.95267 | 365.5613 | [8 empty] | 0.95267 | 368.6157 | 0.95267 | 348.8771 | 0.95267 | 382.5486
3.105 | -136.9117 | 0.95397 | 365.5146 | [8 empty] | 0.95397 | 368.5715 | 0.95397 | 348.8188 | 0.95397 | 382.5169
3.11 | -136.885 | 0.95526 | 365.4668 | [8 empty] | 0.95526 | 368.5261 | 0.95526 | 348.7596 | 0.95526 | 382.4839
3.115 | -136.8582 | 0.95656 | 365.4178 | [8 empty] | 0.95656 | 368.4795 | 0.95656 | 348.6993 | 0.95656 | 382.4496
3.12 | -136.8315 | 0.95785 | 365.3677 | [8 empty] | 0.95785 | 368.4317 | 0.95785 | 348.638 | 0.95785 | 382.414
3.125 | -136.8047 | 0.95915 | 365.3164 | [8 empty] | 0.95915 | 368.3828 | 0.95915 | 348.5757 | 0.95915 | 382.3771
3.13 | -136.682 | 0.96045 | 365.264 | [8 empty] | 0.96045 | 368.3327 | 0.96045 | 348.5124 | 0.96045 | 382.3389
3.135 | -136.6552 | 0.96174 | 365.2104 | [8 empty] | 0.96174 | 368.2814 | 0.96174 | 348.448 | 0.96174 | 382.2994
3.14 | -136.7245 | 0.96304 | 365.1557 | [8 empty] | 0.96304 | 368.2289 | 0.96304 | 348.3827 | 0.96304 | 382.2586
3.145 | -136.7937 | 0.96433 | 365.0998 | [8 empty] | 0.96433 | 368.1753 | 0.96433 | 348.3164 | 0.96433 | 382.2166
3.15 | -136.767 | 0.96563 | 365.0428 | [8 empty] | 0.96563 | 368.1205 | 0.96563 | 348.249 | 0.96563 | 382.1732
3.155 | -136.6442 | 0.96693 | 364.9847 | [8 empty] | 0.96693 | 368.0646 | 0.96693 | 348.1807 | 0.96693 | 382.1286
3.16 | -136.6174 | 0.96822 | 364.9254 | [8 empty] | 0.96822 | 368.0075 | 0.96822 | 348.1114 | 0.96822 | 382.0827
3.165 | -136.5907 | 0.96952 | 364.865 | [8 empty] | 0.96952 | 367.9492 | 0.96952 | 348.0411 | 0.96952 | 382.0355
3.17 | -136.4679 | 0.97081 | 364.8035 | [8 empty] | 0.97081 | 367.8898 | 0.97081 | 347.9697 | 0.97081 | 381.9871
3.175 | -136.4412 | 0.97211 | 364.7409 | [8 empty] | 0.97211 | 367.8292 | 0.97211 | 347.8975 | 0.97211 | 381.9374
3.18 | -136.3184 | 0.97341 | 364.6771 | [8 empty] | 0.97341 | 367.7675 | 0.97341 | 347.8242 | 0.97341 | 381.8864
3.185 | -136.3877 | 0.9747 | 364.6122 | [8 empty] | 0.9747 | 367.7047 | 0.9747 | 347.7499 | 0.9747 | 381.8341
3.19 | -136.3609 | 0.976 | 364.5461 | [8 empty] | 0.976 | 367.6406 | 0.976 | 347.6746 | 0.976 | 381.7806
3.195 | -136.2382 | 0.9773 | 364.479 | [8 empty] | 0.9773 | 367.5755 | 0.9773 | 347.5984 | 0.9773 | 381.7258
3.2 | -136.2114 | 0.97859 | 364.4107 | [8 empty] | 0.97859 | 367.5092 | 0.97859 | 347.5212 | 0.97859 | 381.6697
3.205 | -136.0887 | 0.97989 | 364.3414 | [8 empty] | 0.97989 | 367.4417 | 0.97989 | 347.443 | 0.97989 | 381.6124
3.21 | -135.9659 | 0.98118 | 364.2709 | [8 empty] | 0.98118 | 367.3732 | 0.98118 | 347.3639 | 0.98118 | 381.5538
3.215 | -135.8431 | 0.98248 | 364.1993 | [8 empty] | 0.98248 | 367.3034 | 0.98248 | 347.2838 | 0.98248 | 381.494
3.22 | -135.5284 | 0.98378 | 364.1266 | [8 empty] | 0.98378 | 367.2326 | 0.98378 | 347.2027 | 0.98378 | 381.4329
3.225 | -135.3096 | 0.98507 | 364.0528 | [8 empty] | 0.98507 | 367.1606 | 0.98507 | 347.1206 | 0.98507 | 381.3706
3.23 | -135.0909 | 0.98637 | 363.9779 | [8 empty] | 0.98637 | 367.0875 | 0.98637 | 347.0376 | 0.98637 | 381.307
3.235 | -134.9681 | 0.98766 | 363.9019 | [8 empty] | 0.98766 | 367.0133 | 0.98766 | 346.9536 | 0.98766 | 381.2422
3.24 | -134.9414 | 0.98896 | 363.8247 | [8 empty] | 0.98896 | 366.9379 | 0.98896 | 346.8687 | 0.98896 | 381.1761
3.245 | -134.8186 | 0.99026 | 363.7465 | [8 empty] | 0.99026 | 366.8614 | 0.99026 | 346.7828 | 0.99026 | 381.1087
3.25 | -134.5999 | 0.99155 | 363.6672 | [8 empty] | 0.99155 | 366.7838 | 0.99155 | 346.6959 | 0.99155 | 381.0402
3.255 | -134.3811 | 0.99285 | 363.5869 | [8 empty] | 0.99285 | 366.7051 | 0.99285 | 346.6081 | 0.99285 | 380.9704
3.26 | -134.0664 | 0.99415 | 363.5054 | [8 empty] | 0.99415 | 366.6252 | 0.99415 | 346.5194 | 0.99415 | 380.8993
3.265 | -134.0396 | 0.99544 | 363.4228 | [8 empty] | 0.99544 | 366.5443 | 0.99544 | 346.4297 | 0.99544 | 380.827
3.27 | -133.8209 | 0.99674 | 363.3392 | [8 empty] | 0.99674 | 366.4622 | 0.99674 | 346.339 | 0.99674 | 380.7535
3.275 | -133.6981 | 0.99803 | 363.2544 | [8 empty] | 0.99803 | 366.3791 | 0.99803 | 346.2475 | 0.99803 | 380.6787
3.28 | -133.5753 | 0.99933 | 363.1686 | [8 empty] | 0.99933 | 366.2948 | 0.99933 | 346.1549 | 0.99933 | 380.6027
3.285 | -133.2606 | 1.0006 | 363.0817 | [8 empty] | 1.0006 | 366.2094 | 1.0006 | 346.0615 | 1.0006 | 380.5255
3.29 | -133.0418 | 1.0019 | 362.9937 | [8 empty] | 1.0019 | 366.1229 | 1.0009 | 346.0328 | 1.0019 | 380.4471
3.295 | -132.9191 | 1.0032 | 362.9047 | [8 empty] | 1.0032 | 366.0353 | 1.0011 | 346.0051 | 1.0032 | 380.3674
3.3 | -132.5083 | 1.0045 | 362.8146 | [8 empty] | 1.0045 | 365.9466 | 1.0013 | 345.9769 | 1.0045 | 380.2865
3.305 | -132.3856 | 1.0058 | 362.7234 | [8 empty] | 1.0058 | 365.8569 | 1.0016 | 345.9486 | 1.0058 | 380.2044
3.31 | -132.0708 | 1.0071 | 362.6311 | [8 empty] | 1.0071 | 365.766 | 1.0018 | 345.9199 | 1.0071 | 380.1211
3.315 | -131.7561 | 1.0084 | 362.5378 | [8 empty] | 1.0084 | 365.674 | 1.0021 | 345.8911 | 1.0084 | 380.0366
3.32 | -131.4413 | 1.0097 | 362.4435 | [8 empty] | 1.0097 | 365.581 | 1.0023 | 345.862 | 1.0097 | 379.9508
3.325 | -131.3186 | 1.011 | 362.348 | [8 empty] | 1.011 | 365.4868 | 1.0025 | 345.8326 | 1.011 | 379.8639
3.33 | -131.2918 | 1.0123 | 362.2515 | [8 empty] | 1.0123 | 365.3916 | 1.0028 | 345.803 | 1.0123 | 379.7757
3.335 | -131.169 | 1.0136 | 362.154 | [8 empty] | 1.0136 | 365.2953 | 1.003 | 345.7733 | 1.0136 | 379.6863
3.34 | -130.7583 | 1.0149 | 362.0553 | [8 empty] | 1.0149 | 365.1979 | 1.0033 | 345.7433 | 1.0149 | 379.5957
3.345 | -130.6355 | 1.0162 | 361.9557 | [8 empty] | 1.0162 | 365.0994 | 1.0035 | 345.713 | 1.0162 | 379.5039
3.35 | -130.4168 | 1.0175 | 361.855 | [8 empty] | 1.0175 | 364.9999 | 1.0037 | 345.6825 | 1.0175 | 379.4109
3.355 | -130.198 | 1.0188 | 361.7532 | [8 empty] | 1.0188 | 364.8992 | 1.004 | 345.6517 | 1.0188 | 379.3168
3.36 | -129.9793 | 1.0201 | 361.6504 | [8 empty] | 1.0201 | 364.7975 | 1.0042 | 345.6207 | 1.0201 | 379.2214
3.365 | -129.7605 | 1.0214 | 361.5466 | [8 empty] | 1.0214 | 364.6948 | 1.0045 | 345.5895 | 1.0214 | 379.1248
3.37 | -129.3498 | 1.0227 | 361.4417 | [8 empty] | 1.0227 | 364.591 | 1.0047 | 345.5581 | 1.0227 | 379.027
3.375 | -129.227 | 1.024 | 361.3358 | [8 empty] | 1.024 | 364.4861 | 1.0049 | 345.5264 | 1.024 | 378.928
3.38 | -128.8163 | 1.0253 | 361.2288 | [8 empty] | 1.0253 | 364.3801 | 1.0052 | 345.4945 | 1.0253 | 378.8279
3.385 | -128.5015 | 1.0265 | 361.1208 | [8 empty] | 1.0265 | 364.2731 | 1.0054 | 345.4623 | 1.0265 | 378.7265
3.39 | -128.1868 | 1.0278 | 361.0118 | [8 empty] | 1.0278 | 364.165 | 1.0056 | 345.4298 | 1.0278 | 378.624
3.395 | -127.968 | 1.0291 | 360.9017 | [8 empty] | 1.0291 | 364.0559 | 1.0059 | 345.3971 | 1.0291 | 378.5203
3.4 | -127.7492 | 1.0304 | 360.7906 | [8 empty] | 1.0304 | 363.9457 | 1.0061 | 345.3643 | 1.0304 | 378.4154
3.405 | -127.3385 | 1.0317 | 360.6785 | [8 empty] | 1.0317 | 363.8344 | 1.0064 | 345.3311 | 1.0317 | 378.3093
3.41 | -127.0237 | 1.033 | 360.5654 | [8 empty] | 1.033 | 363.7221 | 1.0066 | 345.2977 | 1.033 | 378.2021
3.415 | -126.901 | 1.0343 | 360.4512 | [8 empty] | 1.0343 | 363.6088 | 1.0068 | 345.264 | 1.0343 | 378.0937
3.42 | -126.5862 | 1.0356 | 360.336 | [8 empty] | 1.0356 | 363.4944 | 1.0071 | 345.2301 | 1.0356 | 377.9841
3.425 | -126.3675 | 1.0369 | 360.2199 | [8 empty] | 1.0369 | 363.379 | 1.0073 | 345.196 | 1.0369 | 377.8733
3.43 | -126.0527 | 1.0382 | 360.1026 | [8 empty] | 1.0382 | 363.2625 | 1.0076 | 345.1616 | 1.0382 | 377.7614
3.435 | -125.738 | 1.0395 | 359.9844 | [8 empty] | 1.0395 | 363.145 | 1.0078 | 345.1269 | 1.0395 | 377.6483
3.44 | -125.5192 | 1.0408 | 359.8652 | [8 empty] | 1.0408 | 363.0264 | 1.008 | 345.092 | 1.0408 | 377.534
3.445 | -125.2045 | 1.0421 | 359.745 | [8 empty] | 1.0421 | 362.9069 | 1.0083 | 345.0569 | 1.0421 | 377.4186
3.45 | -124.7937 | 1.0434 | 359.6237 | [8 empty] | 1.0434 | 362.7862 | 1.0085 | 345.0215 | 1.0434 | 377.302
3.455 | -124.479 | 1.0447 | 359.5015 | [8 empty] | 1.0447 | 362.6646 | 1.0088 | 344.9858 | 1.0447 | 377.1843
3.46 | -124.2602 | 1.046 | 359.3782 | [8 empty] | 1.046 | 362.5419 | 1.009 | 344.9499 | 1.046 | 377.0654
3.465 | -123.9454 | 1.0473 | 359.254 | [8 empty] | 1.0473 | 362.4182 | 1.0092 | 344.9137 | 1.0473 | 376.9453
3.47 | -123.6307 | 1.0486 | 359.1287 | [8 empty] | 1.0486 | 362.2935 | 1.0095 | 344.8773 | 1.0486 | 376.8241
3.475 | -123.2199 | 1.0499 | 359.0025 | [8 empty] | 1.0499 | 362.1677 | 1.0097 | 344.8406 | 1.0499 | 376.7018
3.48 | -123.0012 | 1.0512 | 358.8752 | [8 empty] | 1.0512 | 362.041 | 1.01 | 344.8037 | 1.0512 | 376.5783
3.485 | -122.5904 | 1.0525 | 358.747 | [8 empty] | 1.0525 | 361.9132 | 1.0102 | 344.7664 | 1.0525 | 376.4536
3.49 | -122.1797 | 1.0538 | 358.6178 | [8 empty] | 1.0538 | 361.7844 | 1.0104 | 344.729 | 1.0538 | 376.3278
3.495 | -121.9609 | 1.0551 | 358.4876 | [8 empty] | 1.0551 | 361.6546 | 1.0107 | 344.6913 | 1.0551 | 376.2009
3.5 | -121.6462 | 1.0564 | 358.3564 | [8 empty] | 1.0564 | 361.5237 | 1.0109 | 344.6533 | 1.0564 | 376.0728
3.505 | -121.3314 | 1.0577 | 358.2243 | [8 empty] | 1.0577 | 361.3919 | 1.0111 | 344.615 | 1.0577 | 375.9436
3.51 | -120.9207 | 1.059 | 358.0911 | [8 empty] | 1.059 | 361.259 | 1.0114 | 344.5765 | 1.059 | 375.8133
3.515 | -120.5099 | 1.0602 | 357.957 | [8 empty] | 1.0602 | 361.1252 | 1.0116 | 344.5377 | 1.0602 | 375.6818
3.52 | -120.2911 | 1.0615 | 357.8219 | [8 empty] | 1.0615 | 360.9903 | 1.0119 | 344.4987 | 1.0615 | 375.5492
3.525 | -119.8804 | 1.0628 | 357.6858 | [8 empty] | 1.0628 | 360.8545 | 1.0121 | 344.4594 | 1.0628 | 375.4155
3.53 | -119.5656 | 1.0641 | 357.5488 | [8 empty] | 1.0641 | 360.7176 | 1.0123 | 344.4198 | 1.0641 | 375.2806
3.535 | -119.2509 | 1.0654 | 357.4107 | [8 empty] | 1.0654 | 360.5798 | 1.0126 | 344.38 | 1.0654 | 375.1446
3.54 | -118.9361 | 1.0667 | 357.2718 | [8 empty] | 1.0667 | 360.4409 | 1.0128 | 344.3398 | 1.0667 | 375.0075
3.545 | -118.4294 | 1.068 | 357.1318 | [8 empty] | 1.068 | 360.3011 | 1.0131 | 344.2994 | 1.068 | 374.8693
3.55 | -118.0186 | 1.0693 | 356.9909 | [8 empty] | 1.0693 | 360.1603 | 1.0133 | 344.2588 | 1.0693 | 374.7299
3.555 | -117.7039 | 1.0706 | 356.849 | [8 empty] | 1.0706 | 360.0185 | 1.0135 | 344.2178 | 1.0706 | 374.5895
3.56 | -117.5811 | 1.0719 | 356.7062 | [8 empty] | 1.0719 | 359.8757 | 1.0138 | 344.1766 | 1.0719 | 374.4479
3.565 | -117.2664 | 1.0732 | 356.5624 | [8 empty] | 1.0732 | 359.7319 | 1.014 | 344.1351 | 1.0732 | 374.3052
3.57 | -116.8556 | 1.0745 | 356.4177 | [8 empty] | 1.0745 | 359.5871 | 1.0143 | 344.0934 | 1.0745 | 374.1614
3.575 | -116.4449 | 1.0758 | 356.272 | [8 empty] | 1.0758 | 359.4414 | 1.0145 | 344.0514 | 1.0758 | 374.0165
3.58 | -116.2261 | 1.0771 | 356.1253 | [8 empty] | 1.0771 | 359.2946 | 1.0147 | 344.0091 | 1.0771 | 373.8705
3.585 | -115.8153 | 1.0784 | 355.9777 | [8 empty] | 1.0784 | 359.1469 | 1.015 | 343.9665 | 1.0784 | 373.7233
3.59 | -115.3086 | 1.0797 | 355.8292 | [8 empty] | 1.0797 | 358.9983 | 1.0152 | 343.9236 | 1.0797 | 373.5751
3.595 | -114.9938 | 1.081 | 355.6797 | [8 empty] | 1.081 | 358.8486 | 1.0155 | 343.8805 | 1.081 | 373.4258
3.6 | -114.5831 | 1.0823 | 355.5293 | [8 empty] | 1.0823 | 358.698 | 1.0157 | 343.837 | 1.0823 | 373.2754
3.605 | -114.3643 | 1.0836 | 355.3779 | [8 empty] | 1.0836 | 358.5464 | 1.0159 | 343.7933 | 1.0836 | 373.1238
3.61 | -113.9536 | 1.0849 | 355.2256 | [8 empty] | 1.0849 | 358.3939 | 1.0162 | 343.7493 | 1.0849 | 372.9712
3.615 | -113.4468 | 1.0862 | 355.0724 | [8 empty] | 1.0862 | 358.2404 | 1.0164 | 343.705 | 1.0862 | 372.8175
3.62 | -113.1321 | 1.0875 | 354.9182 | [8 empty] | 1.0875 | 358.0859 | 1.0166 | 343.6605 | 1.0875 | 372.6627
3.625 | -112.7213 | 1.0888 | 354.7631 | [8 empty] | 1.0888 | 357.9305 | 1.0169 | 343.6156 | 1.0888 | 372.5068
3.63 | -112.1186 | 1.0901 | 354.6071 | [8 empty] | 1.0901 | 357.7741 | 1.0171 | 343.5704 | 1.0901 | 372.3499
3.635 | -111.6118 | 1.0914 | 354.4501 | [8 empty] | 1.0914 | 357.6168 | 1.0174 | 343.525 | 1.0914 | 372.1918
3.64 | -111.105 | 1.0927 | 354.2922 | [8 empty] | 1.0927 | 357.4585 | 1.0176 | 343.4793 | 1.0927 | 372.0327
3.645 | -110.6943 | 1.0939 | 354.1334 | [8 empty] | 1.0939 | 357.2993 | 1.0178 | 343.4333 | 1.0939 | 371.8725
3.65 | -110.3795 | 1.0952 | 353.9737 | [8 empty] | 1.0952 | 357.1391 | 1.0181 | 343.387 | 1.0952 | 371.7112
3.655 | -109.9688 | 1.0965 | 353.813 | [8 empty] | 1.0965 | 356.9779 | 1.0183 | 343.3403 | 1.0965 | 371.5488
3.66 | -109.558 | 1.0978 | 353.6515 | [8 empty] | 1.0978 | 356.8159 | 1.0186 | 343.2934 | 1.0978 | 371.3854
3.665 | -109.1473 | 1.0991 | 353.489 | [8 empty] | 1.0991 | 356.6528 | 1.0188 | 343.2462 | 1.0991 | 371.2209
3.67 | -108.7365 | 1.1004 | 353.3256 | [8 empty] | 1.1004 | 356.4889 | 1.019 | 343.1987 | 1.1004 | 371.0553
3.675 | -108.1338 | 1.1017 | 353.1613 | [8 empty] | 1.1017 | 356.324 | 1.0193 | 343.1509 | 1.1017 | 370.8887
3.68 | -107.627 | 1.103 | 352.9961 | [8 empty] | 1.103 | 356.1582 | 1.0195 | 343.1028 | 1.103 | 370.7209
3.685 | -107.1203 | 1.1043 | 352.8299 | [8 empty] | 1.1043 | 355.9914 | 1.0198 | 343.0544 | 1.1043 | 370.5522
3.69 | -106.6135 | 1.1056 | 352.6629 | [8 empty] | 1.1056 | 355.8237 | 1.02 | 343.0057 | 1.1056 | 370.3823
3.695 | -106.2028 | 1.1069 | 352.495 | [8 empty] | 1.1069 | 355.6551 | 1.0202 | 342.9567 | 1.1069 | 370.2114
3.7 | -105.6 | 1.1082 | 352.3262 | [8 empty] | 1.1082 | 355.4855 | 1.0205 | 342.9074 | 1.1082 | 370.0395
3.705 | -105.2852 | 1.1095 | 352.1564 | [8 empty] | 1.1095 | 355.315 | 1.0207 | 342.8578 | 1.1095 | 369.8665
3.71 | -104.9705 | 1.1108 | 351.9858 | [8 empty] | 1.1108 | 355.1436 | 1.021 | 342.8078 | 1.1108 | 369.6924
3.715 | -104.5597 | 1.1121 | 351.8143 | [8 empty] | 1.1121 | 354.9713 | 1.0212 | 342.7576 | 1.1121 | 369.5173
3.72 | -104.053 | 1.1134 | 351.6419 | [8 empty] | 1.1134 | 354.7981 | 1.0214 | 342.707 | 1.1134 | 369.3412
3.725 | -103.6422 | 1.1147 | 351.4686 | [8 empty] | 1.1147 | 354.6239 | 1.0217 | 342.6561 | 1.1147 | 369.164
3.73 | -103.0395 | 1.116 | 351.2944 | [8 empty] | 1.116 | 354.4489 | 1.0219 | 342.605 | 1.116 | 368.9857
3.735 | -102.7247 | 1.1173 | 351.1194 | [8 empty] | 1.1173 | 354.2729 | 1.0221 | 342.5535 | 1.1173 | 368.8064
3.74 | -102.41 | 1.1186 | 350.9434 | [8 empty] | 1.1186 | 354.096 | 1.0224 | 342.5017 | 1.1186 | 368.6261
3.745 | -101.9032 | 1.1199 | 350.7666 | [8 empty] | 1.1199 | 353.9182 | 1.0226 | 342.4495 | 1.1199 | 368.4447
3.75 | -101.4925 | 1.1212 | 350.5889 | [8 empty] | 1.1212 | 353.7395 | 1.0229 | 342.3971 | 1.1212 | 368.2623
3.755 | -100.9857 | 1.1225 | 350.4103 | [8 empty] | 1.1225 | 353.5598 | 1.0231 | 342.3443 | 1.1225 | 368.0789
3.76 | -100.479 | 1.1238 | 350.2308 | [8 empty] | 1.1238 | 353.3793 | 1.0233 | 342.2913 | 1.1238 | 367.8944
3.765 | -100.1642 | 1.1251 | 350.0505 | [8 empty] | 1.1251 | 353.1979 | 1.0236 | 342.2379 | 1.1251 | 367.7089
3.77 | -99.7534 | 1.1264 | 349.8692 | [8 empty] | 1.1264 | 353.0156 | 1.0238 | 342.1842 | 1.1264 | 367.5224
3.775 | -99.3427 | 1.1276 | 349.6872 | [8 empty] | 1.1276 | 352.8324 | 1.0241 | 342.1301 | 1.1276 | 367.3348
3.78 | -98.9319 | 1.1289 | 349.5042 | [8 empty] | 1.1289 | 352.6483 | 1.0243 | 342.0757 | 1.1289 | 367.1462
3.785 | -98.3292 | 1.1302 | 349.3204 | [8 empty] | 1.1302 | 352.4633 | 1.0245 | 342.021 | 1.1302 | 366.9566
3.79 | -97.7264 | 1.1315 | 349.1357 | [8 empty] | 1.1315 | 352.2774 | 1.0248 | 341.966 | 1.1315 | 366.766
3.795 | -97.3157 | 1.1328 | 348.9502 | [8 empty] | 1.1328 | 352.0906 | 1.025 | 341.9106 | 1.1328 | 366.5744
3.8 | -96.7129 | 1.1341 | 348.7638 | [8 empty] | 1.1341 | 351.9029 | 1.0253 | 341.8549 | 1.1341 | 366.3817
3.805 | -96.2062 | 1.1354 | 348.5765 | [8 empty] | 1.1354 | 351.7144 | 1.0255 | 341.7989 | 1.1354 | 366.188
3.81 | -95.6034 | 1.1367 | 348.3884 | [8 empty] | 1.1367 | 351.5249 | 1.0257 | 341.7425 | 1.1367 | 365.9934
3.815 | -95.1927 | 1.138 | 348.1994 | [8 empty] | 1.138 | 351.3346 | 1.026 | 341.6859 | 1.138 | 365.7977
3.82 | -94.6859 | 1.1393 | 348.0096 | [8 empty] | 1.1393 | 351.1434 | 1.0262 | 341.6289 | 1.1393 | 365.601
3.825 | -94.2751 | 1.1406 | 347.819 | [8 empty] | 1.1406 | 350.9514 | 1.0265 | 341.5715 | 1.1406 | 365.4033
3.83 | -93.5764 | 1.1419 | 347.6275 | [8 empty] | 1.1419 | 350.7584 | 1.0267 | 341.5138 | 1.1419 | 365.2046
3.835 | -92.8776 | 1.1432 | 347.4351 | [8 empty] | 1.1432 | 350.5646 | 1.0269 | 341.4557 | 1.1432 | 365.0049
3.84 | -92.2749 | 1.1445 | 347.2419 | [8 empty] | 1.1445 | 350.3699 | 1.0272 | 341.3974 | 1.1445 | 364.8041
3.845 | -91.6721 | 1.1458 | 347.0478 | [8 empty] | 1.1458 | 350.1744 | 1.0274 | 341.3387 | 1.1458 | 364.6024
3.85 | -91.2614 | 1.1471 | 346.853 | [8 empty] | 1.1471 | 349.978 | 1.0276 | 341.2796 | 1.1471 | 364.3997
3.855 | -90.7546 | 1.1484 | 346.6572 | [8 empty] | 1.1484 | 349.7807 | 1.0279 | 341.2202 | 1.1484 | 364.196
3.86 | -90.2479 | 1.1497 | 346.4607 | [8 empty] | 1.1497 | 349.5826 | 1.0281 | 341.1604 | 1.1497 | 363.9913
3.865 | -89.5491 | 1.151 | 346.2633 | [8 empty] | 1.151 | 349.3835 | 1.0284 | 341.1002 | 1.151 | 363.7857
3.87 | -89.0424 | 1.1523 | 346.0651 | [8 empty] | 1.1523 | 349.1837 | 1.0286 | 341.0398 | 1.1523 | 363.579
3.875 | -88.4396 | 1.1536 | 345.866 | [8 empty] | 1.1536 | 348.983 | 1.0288 | 340.979 | 1.1536 | 363.3714
3.88 | -87.8368 | 1.1549 | 345.6662 | [8 empty] | 1.1549 | 348.7814 | 1.0291 | 340.9178 | 1.1549 | 363.1627
3.885 | -87.0421 | 1.1562 | 345.4655 | [8 empty] | 1.1562 | 348.579 | 1.0293 | 340.8562 | 1.1562 | 362.9531
3.89 | -86.3433 | 1.1575 | 345.2639 | [8 empty] | 1.1575 | 348.3757 | 1.0296 | 340.7943 | 1.1575 | 362.7425
3.895 | -85.6446 | 1.1588 | 345.0616 | [8 empty] | 1.1588 | 348.1716 | 1.0298 | 340.732 | 1.1588 | 362.531
3.9 | -85.0418 | 1.1601 | 344.8584 | [8 empty] | 1.1601 | 347.9666 | 1.03 | 340.6694 | 1.1601 | 362.3184
3.905 | -84.3431 | 1.1613 | 344.6544 | [8 empty] | 1.1613 | 347.7608 | 1.0303 | 340.6065 | 1.1613 | 362.1049
3.91 | -83.8363 | 1.1626 | 344.4496 | [8 empty] | 1.1626 | 347.5541 | 1.0305 | 340.5431 | 1.1626 | 361.8904
3.915 | -83.3296 | 1.1639 | 344.244 | [8 empty] | 1.1639 | 347.3467 | 1.0308 | 340.4794 | 1.1639 | 361.6749
3.92 | -82.7268 | 1.1652 | 344.0376 | [8 empty] | 1.1652 | 347.1383 | 1.031 | 340.4153 | 1.1652 | 361.4585
3.925 | -82.2201 | 1.1665 | 343.8304 | [8 empty] | 1.1665 | 346.9292 | 1.0312 | 340.3509 | 1.1665 | 361.2411
3.93 | -81.6173 | 1.1678 | 343.6223 | [8 empty] | 1.1678 | 346.7191 | 1.0315 | 340.2862 | 1.1678 | 361.0228
3.935 | -81.2066 | 1.1691 | 343.4135 | [8 empty] | 1.1691 | 346.5083 | 1.0317 | 340.221 | 1.1691 | 360.8035
3.94 | -80.6998 | 1.1704 | 343.2038 | [8 empty] | 1.1704 | 346.2966 | 1.032 | 340.1555 | 1.1704 | 360.5832
3.945 | -80.0011 | 1.1717 | 342.9934 | [8 empty] | 1.1717 | 346.0842 | 1.0322 | 340.0895 | 1.1717 | 360.3619
3.95 | -79.3983 | 1.173 | 342.7821 | [8 empty] | 1.173 | 345.8708 | 1.0324 | 340.0232 | 1.173 | 360.1398
3.955 | -78.8915 | 1.1743 | 342.5701 | [8 empty] | 1.1743 | 345.6567 | 1.0327 | 339.9566 | 1.1743 | 359.9166
3.96 | -78.3848 | 1.1756 | 342.3572 | [8 empty] | 1.1756 | 345.4417 | 1.0329 | 339.8896 | 1.1756 | 359.6925
3.965 | -77.782 | 1.1769 | 342.1436 | [8 empty] | 1.1769 | 345.2259 | 1.0331 | 339.8222 | 1.1769 | 359.4675
3.97 | -76.9873 | 1.1782 | 341.9291 | [8 empty] | 1.1782 | 345.0093 | 1.0334 | 339.7543 | 1.1782 | 359.2415
3.975 | -76.4805 | 1.1795 | 341.7139 | [8 empty] | 1.1795 | 344.7919 | 1.0336 | 339.6861 | 1.1795 | 359.0146
3.98 | -75.7818 | 1.1808 | 341.4979 | [8 empty] | 1.1808 | 344.5737 | 1.0339 | 339.6175 | 1.1808 | 358.7867
3.985 | -75.083 | 1.1821 | 341.2811 | [8 empty] | 1.1821 | 344.3546 | 1.0341 | 339.5486 | 1.1821 | 358.5579
3.99 | -74.5763 | 1.1834 | 341.0635 | [8 empty] | 1.1834 | 344.1348 | 1.0343 | 339.4792 | 1.1834 | 358.3281
3.995 | -74.0695 | 1.1847 | 340.8452 | [8 empty] | 1.1847 | 343.9141 | 1.0346 | 339.4095 | 1.1847 | 358.0974
4 | -73.3708 | 1.185 | 340.7822 | [8 empty] | 1.186 | 343.6926 | 1.0348 | 339.3394 | 1.186 | 357.8658
4.005 | -72.672 | 1.1853 | 340.7204 | [8 empty] | 1.1873 | 343.4703 | 1.0351 | 339.2688 | 1.1873 | 357.6333
4.01 | -72.0692 | 1.1856 | 340.658 | [8 empty] | 1.1886 | 343.2473 | 1.0353 | 339.1978 | 1.1886 | 357.3998
4.015 | -71.3705 | 1.1859 | 340.5956 | [8 empty] | 1.1899 | 343.0234 | 1.0355 | 339.1265 | 1.1899 | 357.1653
4.02 | -70.6717 | 1.1863 | 340.5329 | [8 empty] | 1.1912 | 342.7987 | 1.0358 | 339.0548 | 1.1912 | 356.93
4.025 | -69.877 | 1.1866 | 340.47 | [8 empty] | 1.1925 | 342.5732 | 1.036 | 338.9827 | 1.1925 | 356.6937
4.03 | -69.3702 | 1.1869 | 340.4068 | [8 empty] | 1.1938 | 342.3469 | 1.0363 | 338.9101 | 1.1938 | 356.4565
4.035 | -68.7675 | 1.1872 | 340.3434 | [8 empty] | 1.195 | 342.1199 | 1.0365 | 338.8372 | 1.195 | 356.2184
4.04 | -68.0687 | 1.1875 | 340.2797 | [8 empty] | 1.1963 | 341.892 | 1.0367 | 338.7638 | 1.1963 | 355.9794
4.045 | -67.37 | 1.1878 | 340.2158 | [8 empty] | 1.1976 | 341.6633 | 1.037 | 338.69 | 1.1976 | 355.7394
4.05 | -66.7672 | 1.1881 | 340.1517 | [8 empty] | 1.1989 | 341.4339 | 1.0372 | 338.6159 | 1.1989 | 355.4986
4.055 | -66.1645 | 1.1885 | 340.0874 | [8 empty] | 1.2002 | 341.2037 | 1.0375 | 338.5413 | 1.2002 | 355.2568
4.06 | -65.5617 | 1.1888 | 340.0228 | [8 empty] | 1.2015 | 340.9727 | 1.0377 | 338.4663 | 1.2015 | 355.0141
4.065 | -64.959 | 1.1891 | 339.9579 | [8 empty] | 1.2028 | 340.7409 | 1.0379 | 338.3909 | 1.2028 | 354.7705
4.07 | -64.3562 | 1.1894 | 339.8928 | [8 empty] | 1.2041 | 340.5083 | 1.0382 | 338.315 | 1.2041 | 354.526
4.075 | -63.6574 | 1.1897 | 339.8275 | [8 empty] | 1.2054 | 340.275 | 1.0384 | 338.2388 | 1.2054 | 354.2806
4.08 | -62.9587 | 1.19 | 339.7619 | [8 empty] | 1.2067 | 340.0408 | 1.0386 | 338.1621 | 1.2067 | 354.0343
4.085 | -62.2599 | 1.1903 | 339.6961 | [8 empty] | 1.208 | 339.806 | 1.0389 | 338.085 | 1.208 | 353.7871
4.09 | -61.6572 | 1.1907 | 339.6301 | [8 empty] | 1.2093 | 339.5703 | 1.0391 | 338.0074 | 1.2093 | 353.5389
4.095 | -61.0544 | 1.191 | 339.5638 | [8 empty] | 1.2106 | 339.3338 | 1.0394 | 337.9294 | 1.2106 | 353.2899
4.1 | -60.4517 | 1.1913 | 339.4972 | [8 empty] | 1.2119 | 339.0966 | 1.0396 | 337.8511 | 1.2119 | 353.04
4.105 | -59.7529 | 1.1916 | 339.4304 | [8 empty] | 1.2132 | 338.8587 | 1.0398 | 337.7723 | 1.2132 | 352.7892
4.11 | -59.1502 | 1.1919 | 339.3634 | [8 empty] | 1.2145 | 338.6199 | 1.0401 | 337.6931 | 1.2145 | 352.5375
4.115 | -58.5474 | 1.1922 | 339.2961 | [8 empty] | 1.2158 | 338.3804 | 1.0403 | 337.6134 | 1.2158 | 352.285
4.12 | -57.8487 | 1.1926 | 339.2286 | [8 empty] | 1.2171 | 338.1402 | 1.0406 | 337.5332 | 1.2171 | 352.0315
4.125 | -57.1499 | 1.1929 | 339.1609 | [8 empty] | 1.2184 | 337.8991 | 1.0408 | 337.4526 | 1.2184 | 351.7771
4.13 | -56.3552 | 1.1932 | 339.0929 | [8 empty] | 1.2197 | 337.6574 | 1.041 | 337.3716 | 1.2197 | 351.5219
4.135 | -55.7524 | 1.1935 | 339.0246 | [8 empty] | 1.221 | 337.4148 | 1.0413 | 337.2902 | 1.221 | 351.2658
4.14 | -55.0536 | 1.1938 | 338.9561 | [8 empty] | 1.2223 | 337.1715 | 1.0415 | 337.2083 | 1.2223 | 351.0088
4.145 | -54.3549 | 1.1941 | 338.8873 | [8 empty] | 1.2236 | 336.9275 | 1.0418 | 337.1259 | 1.2236 | 350.7509
4.15 | -53.5601 | 1.1944 | 338.8183 | [8 empty] | 1.2249 | 336.6827 | 1.042 | 337.043 | 1.2249 | 350.4922
4.155 | -53.0534 | 1.1948 | 338.7491 | [8 empty] | 1.2262 | 336.4372 | 1.0422 | 336.9597 | 1.2262 | 350.2325
4.16 | -52.3546 | 1.1951 | 338.6795 | [8 empty] | 1.2265 | 336.3649 | 1.0425 | 336.876 | 1.2275 | 349.972
4.165 | -51.7519 | 1.1954 | 338.6098 | [8 empty] | 1.2268 | 336.2939 | 1.0427 | 336.7918 | 1.2287 | 349.7107
4.17 | -51.0531 | 1.1957 | 338.5397 | [8 empty] | 1.2272 | 336.2223 | 1.043 | 336.7072 | 1.23 | 349.4484
4.175 | -50.5464 | 1.196 | 338.4694 | [8 empty] | 1.2275 | 336.1506 | 1.0432 | 336.622 | 1.2313 | 349.1854
|  | 1.1963 | 338.3989 | [8 empty] | 1.2278 | 336.0787 | 1.0434 | 336.5364 | 1.2326 | 348.9214
|  | 1.1966 | 338.3282 | [8 empty] | 1.2282 | 336.0065 | 1.0437 | 336.4503 | 1.2339 | 348.6566
|  | 1.197 | 338.2571 | [8 empty] | 1.2285 | 335.9341 | 1.0439 | 336.3638 | 1.2352 | 348.3909
|  | 1.1973 | 338.1858 | [8 empty] | 1.2288 | 335.8615 | 1.0441 | 336.2768 | 1.2365 | 348.1243
|  | 1.1976 | 338.1143 | [8 empty] | 1.2291 | 335.7886 | 1.0444 | 336.1893 | 1.2378 | 347.8569
|  | 1.1979 | 338.0424 | [8 empty] | 1.2295 | 335.7155 | 1.0446 | 336.1013 | 1.2391 | 347.5887
|  | 1.1982 | 337.9704 | [8 empty] | 1.2298 | 335.6422 | 1.0449 | 336.0128 | 1.2404 | 347.3196
|  | 1.1985 | 337.8981 | [8 empty] | 1.2301 | 335.5686 | 1.0451 | 335.924 | 1.2417 | 347.0496
|  | 1.1988 | 337.8255 | [8 empty] | 1.2305 | 335.4948 | 1.0453 | 335.8346 | 1.243 | 346.7788
|  | 1.1992 | 337.7526 | [8 empty] | 1.2308 | 335.4208 | 1.0456 | 335.7447 | 1.2443 | 346.5072
|  | 1.1995 | 337.6795 | [8 empty] | 1.2311 | 335.3465 | 1.0458 | 335.6543 | 1.2456 | 346.2347
|  | 1.1998 | 337.6061 | [8 empty] | 1.2315 | 335.272 | 1.0461 | 335.5634 | 1.2469 | 345.9614
|  | 1.2001 | 337.5324 | [8 empty] | 1.2318 | 335.1972 | 1.0463 | 335.4719 | 1.2482 | 345.6872
|  | 1.2004 | 337.4585 | [8 empty] | 1.2321 | 335.1223 | 1.0465 | 335.3801 | 1.2495 | 345.4122
|  | 1.2007 | 337.3844 | [8 empty] | 1.2325 | 335.0471 | 1.0468 | 335.2878 | 1.2508 | 345.1363
|  | 1.2011 | 337.3099 | [8 empty] | 1.2328 | 334.9716 | 1.047 | 335.1949 | 1.2521 | 344.8597
|  | 1.2014 | 337.2352 | [8 empty] | 1.2331 | 334.8959 | 1.0473 | 335.1015 | 1.2534 | 344.5821
|  | 1.2017 | 337.1602 | [8 empty] | 1.2335 | 334.82 | 1.0475 | 335.0075 | 1.2547 | 344.3038
|  | 1.202 | 337.085 | [8 empty] | 1.2338 | 334.7438 | 1.0477 | 334.913 | 1.256 | 344.0246
|  | 1.2023 | 337.0095 | [8 empty] | 1.2341 | 334.6674 | 1.048 | 334.8182 | 1.2573 | 343.7446
|  | 1.2026 | 336.9337 | [8 empty] | 1.2345 | 334.5908 | 1.0482 | 334.7228 | 1.2586 | 343.4638
|  | 1.2029 | 336.8577 | [8 empty] | 1.2348 | 334.5139 | 1.0485 | 334.6268 | 1.2599 | 343.1821
|  | 1.2033 | 336.7813 | [8 empty] | 1.2351 | 334.4368 | 1.0487 | 334.5303 | 1.2612 | 342.8997
|  | 1.2036 | 336.7047 | [8 empty] | 1.2355 | 334.3594 | 1.0489 | 334.4332 | 1.2624 | 342.6164
|  | 1.2039 | 336.6278 | [8 empty] | 1.2358 | 334.2818 | 1.0492 | 334.3355 | 1.2637 | 342.3323
|  | 1.2042 | 336.5507 | [8 empty] | 1.2361 | 334.2039 | 1.0494 | 334.2375 | 1.265 | 342.0474
|  | 1.2045 | 336.4733 | [8 empty] | 1.2365 | 334.1259 | 1.0496 | 334.1389 | 1.2663 | 341.7616
|  | 1.2048 | 336.3956 | [8 empty] | 1.2368 | 334.0475 | 1.0499 | 334.0398 | 1.2676 | 341.4751
|  | 1.2051 | 336.3177 | [8 empty] | 1.2371 | 333.969 | 1.0501 | 333.9401 | 1.2689 | 341.1877
|  | 1.2055 | 336.2394 | [8 empty] | 1.2375 | 333.8901 | 1.0504 | 333.8397 | 1.2702 | 340.8996
|  | 1.2058 | 336.1608 | [8 empty] | 1.2378 | 333.811 | 1.0506 | 333.739 | 1.2715 | 340.6106
|  | 1.2061 | 336.082 | [8 empty] | 1.2381 | 333.7317 | 1.0508 | 333.6377 | 1.2728 | 340.3208
|  | 1.2064 | 336.0029 | [8 empty] | 1.2385 | 333.6521 | 1.0511 | 333.5358 | 1.2741 | 340.0302
|  | 1.2067 | 335.9236 | [8 empty] | 1.2388 | 333.5723 | 1.0513 | 333.4334 | 1.2754 | 339.7389
|  | 1.207 | 335.8439 | [8 empty] | 1.2391 | 333.4923 | 1.0516 | 333.3303 | 1.2767 | 339.4467
|  | 1.2073 | 335.764 | [8 empty] | 1.2395 | 333.412 | 1.0518 | 333.2267 | 1.278 | 339.1537
|  | 1.2077 | 335.6837 | [8 empty] | 1.2398 | 333.3314 | 1.052 | 333.1226 | 1.2793 | 338.86
|  | 1.208 | 335.6032 | [8 empty] | 1.2401 | 333.2505 | 1.0523 | 333.018 | 1.2806 | 338.5654
|  | 1.2083 | 335.5224 | [8 empty] | 1.2405 | 333.1694 | 1.0525 | 332.9127 | 1.2819 | 338.2701
|  | 1.2086 | 335.4414 | [8 empty] | 1.2408 | 333.0881 | 1.0528 | 332.8068 | 1.2832 | 337.9739
|  | 1.2089 | 335.36 | [8 empty] | 1.2411 | 333.0066 | 1.053 | 332.7003 | 1.2845 | 337.677
|  | 1.2092 | 335.2784 | [8 empty] | 1.2415 | 332.9247 | 1.0532 | 332.5932 | 1.2858 | 337.3793
|  | 1.2096 | 335.1964 | [8 empty] | 1.2418 | 332.8427 | 1.0535 | 332.4857 | 1.2871 | 337.0808
|  | 1.2099 | 335.1141 | [8 empty] | 1.2421 | 332.7603 | 1.0537 | 332.3775 | 1.2884 | 336.7815
|  | 1.2102 | 335.0316 | [8 empty] | 1.2424 | 332.6777 | 1.054 | 332.2687 | 1.2897 | 336.4815
|  | 1.2105 | 334.9488 | [8 empty] | 1.2428 | 332.5948 | 1.0542 | 332.1593 | 1.291 | 336.1807
|  | 1.2108 | 334.8657 | [8 empty] | 1.2431 | 332.5118 | 1.0544 | 332.0493 | 1.2923 | 335.8791
|  | 1.2111 | 334.7823 | [8 empty] | 1.2434 | 332.4284 | 1.0547 | 331.9386 | 1.2936 | 335.5767
|  | 1.2114 | 334.6986 | [8 empty] | 1.2438 | 332.3448 | 1.0549 | 331.8275 | 1.2949 | 335.2735
|  | 1.2118 | 334.6146 | [8 empty] | 1.2441 | 332.2609 | 1.0551 | 331.7158 | 1.2961 | 334.9696
|  | 1.2121 | 334.5303 | [8 empty] | 1.2444 | 332.1767 | 1.0554 | 331.6033 | 1.2974 | 334.6649
|  | 1.2124 | 334.4458 | [8 empty] | 1.2448 | 332.0923 | 1.0556 | 331.4903 | 1.2987 | 334.3595
|  | 1.2127 | 334.3609 | [8 empty] | 1.2451 | 332.0076 | 1.0559 | 331.3765 | 1.3 | 334.0533
|  | 1.213 | 334.2757 | [8 empty] | 1.2454 | 331.9227 | 1.0561 | 331.2621 | 1.3013 | 333.7463
|  | 1.2133 | 334.1902 | [8 empty] | 1.2458 | 331.8375 | 1.0563 | 331.1473 | 1.3026 | 333.4386
|  | 1.2136 | 334.1044 | [8 empty] | 1.2461 | 331.7521 | 1.0566 | 331.0318 | 1.3039 | 333.1301
|  | 1.214 | 334.0183 | [8 empty] | 1.2464 | 331.6663 | 1.0568 | 330.9156 | 1.3052 | 332.8208
|  | 1.2143 | 333.932 | [8 empty] | 1.2468 | 331.5803 | 1.0571 | 330.7988 | 1.3065 | 332.5108
|  | 1.2146 | 333.8453 | [8 empty] | 1.2471 | 331.4941 | 1.0573 | 330.6812 | 1.3078 | 332.2001
|  | 1.2149 | 333.7583 | [8 empty] | 1.2474 | 331.4076 | 1.0575 | 330.5632 | 1.3091 | 331.8885
|  | 1.2152 | 333.671 | [8 empty] | 1.2478 | 331.3208 | 1.0578 | 330.4445 | 1.3104 | 331.5763
|  | 1.2155 | 333.5833 | [8 empty] | 1.2481 | 331.2337 | 1.058 | 330.3251 | 1.3117 | 331.2633
|  | 1.2158 | 333.4953 | [8 empty] | 1.2484 | 331.1464 | 1.0583 | 330.205 | 1.313 | 330.9495
|  | 1.2162 | 333.4071 | [8 empty] | 1.2488 | 331.0588 | 1.0585 | 330.0842 | 1.3143 | 330.635
|  | 1.2165 | 333.3186 | [8 empty] | 1.2491 | 330.971 | 1.0587 | 329.9627 | 1.3156 | 330.3198
|  | 1.2168 | 333.2297 | [8 empty] | 1.2494 | 330.8829 | 1.059 | 329.8407 | 1.3169 | 330.0038
|  | 1.2171 | 333.1405 | [8 empty] | 1.2498 | 330.7945 | 1.0592 | 329.718 | 1.3182 | 329.6871
|  | 1.2174 | 333.051 | [8 empty] | 1.2501 | 330.7058 | 1.0595 | 329.5945 | 1.3195 | 329.3697
|  | 1.2177 | 332.9612 | [8 empty] | 1.2504 | 330.6168 | 1.0597 | 329.4704 | 1.3208 | 329.0515
|  | 1.2181 | 332.8711 | [8 empty] | 1.2508 | 330.5276 | 1.0599 | 329.3455 | 1.3221 | 328.7326
|  | 1.2184 | 332.7807 | [8 empty] | 1.2511 | 330.438 | 1.0602 | 329.2199 | 1.3234 | 328.4129
|  | 1.2187 | 332.6899 | [8 empty] | 1.2514 | 330.3483 | 1.0604 | 329.0938 | 1.3247 | 328.0925
|  | 1.219 | 332.5988 | [8 empty] | 1.2518 | 330.2583 | 1.0606 | 328.9669 | 1.326 | 327.7714
|  | 1.2193 | 332.5074 | [8 empty] | 1.2521 | 330.1679 | 1.0609 | 328.8393 | 1.3273 | 327.4496
|  | 1.2196 | 332.4157 | [8 empty] | 1.2524 | 330.0773 | 1.0611 | 328.711 | 1.3286 | 327.1271
|  | 1.2199 | 332.3236 | [8 empty] | 1.2528 | 329.9864 | 1.0614 | 328.5819 | 1.3298 | 326.8038
|  | 1.2203 | 332.2313 | [8 empty] | 1.2531 | 329.8952 | 1.0616 | 328.452 | 1.3311 | 326.4798
|  | 1.2206 | 332.1386 | [8 empty] | 1.2534 | 329.8038 | 1.0618 | 328.3216 | 1.3324 | 326.1551
|  | 1.2209 | 332.0456 | [8 empty] | 1.2538 | 329.7121 | 1.0621 | 328.1904 | 1.3337 | 325.8297
|  | 1.2212 | 331.9522 | [8 empty] | 1.2541 | 329.6201 | 1.0623 | 328.0584 | 1.335 | 325.5035
|  | 1.2215 | 331.8585 | [8 empty] | 1.2544 | 329.5278 | 1.0626 | 327.9257 | 1.3363 | 325.1767
|  | 1.2218 | 331.7645 | [8 empty] | 1.2548 | 329.4352 | 1.0628 | 327.7922 | 1.3376 | 324.8491
|  | 1.2221 | 331.6702 | [8 empty] | 1.2551 | 329.3423 | 1.063 | 327.6579 | 1.3389 | 324.5209
|  | 1.2225 | 331.5755 | [8 empty] | 1.2554 | 329.2492 | 1.0633 | 327.523 | 1.3402 | 324.1919
|  | 1.2228 | 331.4805 | [8 empty] | 1.2557 | 329.1558 | 1.0635 | 327.3874 | 1.3415 | 323.8622
|  | 1.2231 | 331.3852 | [8 empty] | 1.2561 | 329.062 | 1.0638 | 327.2509 | 1.3428 | 323.5318
|  | 1.2234 | 331.2894 | [8 empty] | 1.2564 | 328.968 | 1.064 | 327.1135 | 1.3441 | 323.2007
|  | 1.2237 | 331.1935 | [8 empty] | 1.2567 | 328.8737 | 1.0642 | 326.9754 | 1.3454 | 322.869
|  | 1.224 | 331.0971 | [8 empty] | 1.2571 | 328.7791 | 1.0645 | 326.8365 | 1.3467 | 322.5365
|  | 1.2243 | 331.0004 | [8 empty] | 1.2574 | 328.6843 | 1.0647 | 326.697 | 1.348 | 322.2033
|  | 1.2247 | 330.9034 | [8 empty] | 1.2577 | 328.5891 | 1.065 | 326.5567 | 1.3493 | 321.8694
|  | 1.225 | 330.806 | [8 empty] | 1.2581 | 328.4936 | 1.0652 | 326.4155 | 1.3506 | 321.5349
|  | 1.2253 | 330.7082 | [8 empty] | 1.2584 | 328.3979 | 1.0654 | 326.2734 | 1.3519 | 321.1996
|  | 1.2256 | 330.6102 | [8 empty] | 1.2587 | 328.3018 | 1.0657 | 326.1305 | 1.3532 | 320.8637
|  | 1.2259 | 330.5118 | [8 empty] | 1.2591 | 328.2054 | 1.0659 | 325.9867 | 1.3545 | 320.5271
|  | 1.2262 | 330.4131 | [8 empty] | 1.2594 | 328.1088 | 1.0661 | 325.8421 | 1.3558 | 320.1898
|  | 1.2266 | 330.314 | [8 empty] | 1.2597 | 328.0118 | 1.0664 | 325.6969 | 1.3571 | 319.8518
|  | 1.2269 | 330.2145 | [8 empty] | 1.2601 | 327.9146 | 1.0666 | 325.5508 | 1.3584 | 319.5131
|  | 1.2272 | 330.1146 | [8 empty] | 1.2604 | 327.8171 | 1.0669 | 325.4039 | 1.3597 | 319.1737
|  | 1.2275 | 330.0145 | [8 empty] | 1.2607 | 327.7192 | 1.0671 | 325.256 | 1.361 | 318.8337
|  | 1.2278 | 329.914 | [8 empty] | 1.2611 | 327.6211 | 1.0673 | 325.1071 | 1.3623 | 318.493
|  | 1.2281 | 329.8132 | [8 empty] | 1.2614 | 327.5226 | 1.0676 | 324.9574 | 1.3635 | 318.1516
|  | 1.2284 | 329.7119 | [8 empty] | 1.2617 | 327.4239 | 1.0678 | 324.807 | 1.3648 | 317.8096
|  | 1.2288 | 329.6103 | [8 empty] | 1.2621 | 327.3248 | 1.0681 | 324.6559 | 1.3661 | 317.4668
|  | 1.2291 | 329.5083 | [8 empty] | 1.2624 | 327.2255 | 1.0683 | 324.5037 | 1.3674 | 317.1234
|  | 1.2294 | 329.4061 | [8 empty] | 1.2627 | 327.1258 | 1.0685 | 324.3507 | 1.3687 | 316.7794
|  | 1.2297 | 329.3034 | [8 empty] | 1.2631 | 327.0258 | 1.0688 | 324.1966 | 1.37 | 316.4347
|  | 1.23 | 329.2004 | [8 empty] | 1.2634 | 326.9255 | 1.069 | 324.0416 | 1.3713 | 316.0893
|  | 1.2303 | 329.097 | [8 empty] | 1.2637 | 326.8249 | 1.0693 | 323.886 | 1.3726 | 315.7432
|  | 1.2306 | 328.9932 | [8 empty] | 1.2641 | 326.7241 | 1.0695 | 323.7295 | 1.3739 | 315.3965
|  | 1.231 | 328.889 | [8 empty] | 1.2644 | 326.6229 | 1.0697 | 323.572 | 1.3752 | 315.0491
|  | 1.2313 | 328.7846 | [8 empty] | 1.2647 | 326.5213 | 1.07 | 323.4135 | 1.3765 | 314.7011
|  | 1.2316 | 328.6797 | [8 empty] | 1.2651 | 326.4195 | 1.0702 | 323.254 | 1.3778 | 314.3524
|  | 1.2319 | 328.5745 | [8 empty] | 1.2654 | 326.3173 | 1.0705 | 323.0935 | 1.3791 | 314.0031
|  | 1.2322 | 328.4689 | [8 empty] | 1.2657 | 326.2149 | 1.0707 | 322.932 | 1.3804 | 313.6531
|  | 1.2325 | 328.3628 | [8 empty] | 1.2661 | 326.1121 | 1.0709 | 322.7698 | 1.3817 | 313.3025
|  | 1.2328 | 328.2564 | [8 empty] | 1.2664 | 326.009 | 1.0712 | 322.6067 | 1.383 | 312.9512
|  | 1.2332 | 328.1497 | [8 empty] | 1.2667 | 325.9056 | 1.0714 | 322.4425 | 1.3843 | 312.5993
|  | 1.2335 | 328.0426 | [8 empty] | 1.2671 | 325.8018 | 1.0716 | 322.2773 | 1.3856 | 312.2467
|  | 1.2338 | 327.9351 | [8 empty] | 1.2674 | 325.6977 | 1.0719 | 322.1111 | 1.3869 | 311.8935
|  | 1.2341 | 327.8272 | [8 empty] | 1.2677 | 325.5934 | 1.0721 | 321.9438 | 1.3882 | 311.5397
|  | 1.2344 | 327.7189 | [8 empty] | 1.2681 | 325.4887 | 1.0724 | 321.7754 | 1.3895 | 311.1852
|  | 1.2347 | 327.6102 | [8 empty] | 1.2684 | 325.3837 | 1.0726 | 321.6063 | 1.3908 | 310.8301
|  | 1.2351 | 327.5012 | [8 empty] | 1.2687 | 325.2783 | 1.0728 | 321.4362 | 1.3921 | 310.4743
|  | 1.2354 | 327.3917 | [8 empty] | 1.269 | 325.1727 | 1.0731 | 321.2651 | 1.3934 | 310.1179
|  | 1.2357 | 327.2819 | [8 empty] | 1.2694 | 325.0666 | 1.0733 | 321.0928 | 1.3947 | 309.7609
|  | 1.236 | 327.1717 | [8 empty] | 1.2697 | 324.9603 | 1.0736 | 320.9194 | 1.396 | 309.4033
|  | 1.2363 | 327.061 | [8 empty] | 1.27 | 324.8537 | 1.0738 | 320.7449 | 1.3972 | 309.045
|  | 1.2366 | 326.9501 | [8 empty] | 1.2704 | 324.7467 | 1.074 | 320.5697 | 1.3985 | 308.6861
|  | 1.2369 | 326.8387 | [8 empty] | 1.2707 | 324.6394 | 1.0743 | 320.3933 | 1.3998 | 308.3266
|  | 1.2373 | 326.7269 | [8 empty] | 1.271 | 324.5317 | 1.0745 | 320.2159 | 1.4011 | 307.9664
|  | 1.2376 | 326.6147 | [8 empty] | 1.2714 | 324.4237 | 1.0748 | 320.0373 | 1.4024 | 307.6056
|  | 1.2379 | 326.5021 | [8 empty] | 1.2717 | 324.3154 | 1.075 | 319.8576 | 1.4037 | 307.2442
|  | 1.2382 | 326.389 | [8 empty] | 1.272 | 324.2068 | 1.0752 | 319.6767 | 1.405 | 306.8822
|  | 1.2385 | 326.2757 | [8 empty] | 1.2724 | 324.0978 | 1.0755 | 319.4946 | 1.4063 | 306.5196
|  | 1.2388 | 326.1619 | [8 empty] | 1.2727 | 323.9884 | 1.0757 | 319.3113 | 1.4076 | 306.1564
|  | 1.2391 | 326.0477 | [8 empty] | 1.273 | 323.8787 | 1.076 | 319.1272 | 1.4089 | 305.7925
|  | 1.2395 | 325.933 | [8 empty] | 1.2734 | 323.7687 | 1.0762 | 318.942 | 1.4102 | 305.4281
|  | 1.2398 | 325.8179 | [8 empty] | 1.2737 | 323.6584 | 1.0764 | 318.7556 | 1.4115 | 305.063
|  | 1.2401 | 325.7024 | [8 empty] | 1.274 | 323.5477 | 1.0767 | 318.568 | 1.4128 | 304.6973
|  | 1.2404 | 325.5865 | [8 empty] | 1.2744 | 323.4367 | 1.0769 | 318.3791 | 1.4141 | 304.331
|  | 1.2407 | 325.4703 | [8 empty] | 1.2747 | 323.3253 | 1.0771 | 318.1889 | 1.4154 | 303.9641
|  | 1.241 | 325.3536 | [8 empty] | 1.275 | 323.2135 | 1.0774 | 317.9975 | 1.4167 | 303.5967
|  | 1.2413 | 325.2364 | [8 empty] | 1.2754 | 323.1014 | 1.0776 | 317.8053 | 1.418 | 303.2286
|  | 1.2417 | 325.1188 | [8 empty] | 1.2757 | 322.9889 | 1.0779 | 317.6119 | 1.4193 | 302.8599
|  | 1.242 | 325.0007 | [8 empty] | 1.276 | 322.8762 | 1.0781 | 317.4172 | 1.4206 | 302.4906
|  | 1.2423 | 324.8823 | [8 empty] | 1.2764 | 322.7631 | 1.0783 | 317.2212 | 1.4219 | 302.1207
|  | 1.2426 | 324.7635 | [8 empty] | 1.2767 | 322.6496 | 1.0786 | 317.0238 | 1.4232 | 301.7503
|  | 1.2429 | 324.6442 | [8 empty] | 1.277 | 322.5357 | 1.0788 | 316.8252 | 1.4245 | 301.3792
|  | 1.2432 | 324.5245 | [8 empty] | 1.2774 | 322.4214 | 1.0791 | 316.6251 | 1.4258 | 301.0075
|  | 1.2436 | 324.4043 | [8 empty] | 1.2777 | 322.3069 | 1.0793 | 316.4238 | 1.4271 | 300.6353
|  | 1.2439 | 324.2836 | [8 empty] | 1.278 | 322.192 | 1.0795 | 316.2215 | 1.4284 | 300.2625
|  | 1.2442 | 324.1627 | [8 empty] | 1.2784 | 322.0767 | 1.0798 | 316.018 | 1.4297 | 299.8891
|  | 1.2445 | 324.0412 | [8 empty] | 1.2787 | 321.9611 | 1.08 | 315.8131 | 1.4309 | 299.5151
|  | 1.2448 | 323.9193 | [8 empty] | 1.279 | 321.845 | 1.0803 | 315.6068 | 1.4322 | 299.1405
|  | 1.2451 | 323.7969 | [8 empty] | 1.2794 | 321.7286 | 1.0805 | 315.3991 | 1.4335 | 298.7654
|  | 1.2454 | 323.674 | [8 empty] | 1.2797 | 321.6119 | 1.0807 | 315.19 | 1.4348 | 298.3896
|  | 1.2458 | 323.5507 | [8 empty] | 1.28 | 321.4948 | 1.081 | 314.9794 | 1.4361 | 298.0133
|  | 1.2461 | 323.427 | [8 empty] | 1.2804 | 321.3773 | 1.0812 | 314.7675 | 1.4374 | 297.6364
|  | 1.2464 | 323.3029 | [8 empty] | 1.2807 | 321.2595 | 1.0815 | 314.5542 | 1.4387 | 297.259
|  | 1.2467 | 323.1783 | [8 empty] | 1.281 | 321.1412 | 1.0817 | 314.3399 | 1.44 | 296.881
|  | 1.247 | 323.0532 | [8 empty] | 1.2814 | 321.0226 | 1.0819 | 314.1241 | 1.4413 | 296.5024
|  | 1.2473 | 322.9276 | [8 empty] | 1.2817 | 320.9037 | 1.0822 | 313.9068 | 1.4426 | 296.1232
|  | 1.2476 | 322.8017 | [8 empty] | 1.282 | 320.7844 | 1.0824 | 313.6879 | 1.4439 | 295.7435
|  | 1.248 | 322.6753 | [8 empty] | 1.2823 | 320.6646 | 1.0826 | 313.4675 | 1.4452 | 295.3632
|  | 1.2483 | 322.5488 | [8 empty] | 1.2827 | 320.5445 | 1.0829 | 313.2458 | 1.4465 | 294.9823
|  | 1.2486 | 322.4218 | [8 empty] | 1.283 | 320.424 | 1.0831 | 313.0226 | 1.4478 | 294.6009
|  | 1.2489 | 322.294 | [8 empty] | 1.2833 | 320.3031 | 1.0834 | 312.7977 | 1.4491 | 294.219
|  | 1.2492 | 322.1659 | [8 empty] | 1.2837 | 320.1819 | 1.0836 | 312.5713 | 1.4504 | 293.8364
|  | 1.2495 | 322.0373 | [8 empty] | 1.284 | 320.0603 | 1.0838 | 312.3432 | 1.4517 | 293.4533
|  | 1.2498 | 321.9082 | [8 empty] | 1.2843 | 319.9383 | 1.0841 | 312.1141 | 1.453 | 293.0697
|  | 1.2502 | 321.7786 | [8 empty] | 1.2847 | 319.8158 | 1.0843 | 311.8833 | 1.4543 | 292.6855
|  | 1.2505 | 321.6485 | [8 empty] | 1.285 | 319.693 | 1.0846 | 311.6509 | 1.4556 | 292.3008
|  | 1.2508 | 321.518 | [8 empty] | 1.2853 | 319.5698 | 1.0848 | 311.417 | 1.4569 | 291.9155
|  | 1.2511 | 321.387 | [8 empty] | 1.2857 | 319.4462 | 1.085 | 311.1815 | 1.4582 | 291.5296
|  | 1.2514 | 321.2555 | [8 empty] | 1.286 | 319.3223 | 1.0853 | 310.9443 | 1.4595 | 291.1432
|  | 1.2517 | 321.1234 | [8 empty] | 1.2863 | 319.1979 | 1.0855 | 310.7053 | 1.4608 | 290.7563
|  | 1.2521 | 320.9909 | [8 empty] | 1.2867 | 319.0731 | 1.0858 | 310.4647 | 1.4621 | 290.3688
|  | 1.2524 | 320.8579 | [8 empty] | 1.287 | 318.9479 | 1.086 | 310.2223 | 1.4634 | 289.9808
|  | 1.2527 | 320.7245 | [8 empty] | 1.2873 | 318.8223 | 1.0862 | 309.9781 | 1.4646 | 289.5923
|  | 1.253 | 320.5905 | [8 empty] | 1.2877 | 318.6963 | 1.0865 | 309.7324 | 1.4659 | 289.2032
|  | 1.2533 | 320.4559 | [8 empty] | 1.288 | 318.57 | 1.0867 | 309.485 | 1.4672 | 288.8136
|  | 1.2536 | 320.3209 | [8 empty] | 1.2883 | 318.4432 | 1.087 | 309.2364 | 1.4685 | 288.4234
|  | 1.2539 | 320.1855 | [8 empty] | 1.2887 | 318.316 | 1.0872 | 308.9859 | 1.4698 | 288.0327
|  | 1.2543 | 320.0495 | [8 empty] | 1.289 | 318.1883 | 1.0874 | 308.7336 | 1.4711 | 287.6415
|  | 1.2546 | 319.913 | [8 empty] | 1.2893 | 318.0604 | 1.0877 | 308.4794 | 1.4724 | 287.2497
|  | 1.2549 | 319.7759 | [8 empty] | 1.2897 | 317.932 | 1.0879 | 308.2233 | 1.4737 | 286.8575
|  | 1.2552 | 319.6382 | [8 empty] | 1.29 | 317.8031 | 1.0881 | 307.9657 | 1.4742 | 286.7216
|  | 1.2555 | 319.5002 | [8 empty] | 1.2903 | 317.6739 | 1.0884 | 307.7062 | 1.4746 | 286.5867
|  | 1.2558 | 319.3616 | [8 empty] | 1.2907 | 317.5442 | 1.0886 | 307.4447 | 1.475 | 286.4514
|  | 1.2561 | 319.2225 | [8 empty] | 1.291 | 317.414 | 1.0889 | 307.1814 | 1.4755 | 286.3161
|  | 1.2565 | 319.0828 | [8 empty] | 1.2913 | 317.2836 | 1.0891 | 306.9161 | 1.4759 | 286.1806
|  | 1.2568 | 318.9426 | [8 empty] | 1.2917 | 317.1527 | 1.0893 | 306.6488 | 1.4763 | 286.0449
|  | 1.2571 | 318.8017 | [8 empty] | 1.292 | 317.0213 | 1.0896 | 306.3796 | 1.4768 | 285.909
|  | 1.2574 | 318.6605 | [8 empty] | 1.2923 | 316.8895 | 1.0898 | 306.1084 | 1.4772 | 285.773
|  | 1.2577 | 318.5187 | [8 empty] | 1.2927 | 316.7572 | 1.0901 | 305.8355 | 1.4776 | 285.6368
|  | 1.258 | 318.3763 | [8 empty] | 1.293 | 316.6245 | 1.0903 | 305.5606 | 1.4781 | 285.5004
|  | 1.2583 | 318.2333 | [8 empty] | 1.2933 | 316.4915 | 1.0905 | 305.2836 | 1.4785 | 285.3638
|  | 1.2587 | 318.0898 | [8 empty] | 1.2937 | 316.358 | 1.0908 | 305.0047 | 1.479 | 285.2271
|  | 1.259 | 317.9458 | [8 empty] | 1.294 | 316.2241 | 1.091 | 304.7236 | 1.4794 | 285.0903
|  | 1.2593 | 317.8013 | [8 empty] | 1.2943 | 316.0896 | 1.0913 | 304.4405 | 1.4798 | 284.9532
|  | 1.2596 | 317.6562 | [8 empty] | 1.2947 | 315.9547 | 1.0915 | 304.1553 | 1.4803 | 284.816
|  | 1.2599 | 317.5104 | [8 empty] | 1.295 | 315.8194 | 1.0917 | 303.8684 | 1.4807 | 284.6786
|  | 1.2602 | 317.3641 | [8 empty] | 1.2953 | 315.6837 | 1.092 | 303.5793 | 1.4811 | 284.5411
|  | 1.2606 | 317.2173 | [8 empty] | 1.2956 | 315.5476 | 1.0922 | 303.2882 | 1.4816 | 284.4033
|  | 1.2609 | 317.07 | [8 empty] | 1.296 | 315.411 | 1.0925 | 302.9949 | 1.482 | 284.2654
|  | 1.2612 | 316.922 | [8 empty] | 1.2963 | 315.2738 | 1.0927 | 302.6995 | 1.4825 | 284.1274
|  | 1.2615 | 316.7734 | [8 empty] | 1.2966 | 315.1363 | 1.0929 | 302.4048 | 1.4829 | 283.9892
|  | 1.2618 | 316.624 | [8 empty] | 1.297 | 314.9983 | 1.0932 | 302.1067 | 1.4833 | 283.8508
|  | 1.2621 | 316.4742 | [8 empty] | 1.2973 | 314.8599 | 1.0934 | 301.8051 | 1.4838 | 283.7122
|  | 1.2624 | 316.3238 | [8 empty] | 1.2976 | 314.721 | 1.0936 | 301.5009 | 1.4842 | 283.5735
|  | 1.2628 | 316.1728 | [8 empty] | 1.298 | 314.5817 | 1.0939 | 301.1949 | 1.4846 | 283.4346
|  | 1.2631 | 316.0212 | [8 empty] | 1.2983 | 314.4418 | 1.0941 | 300.8862 | 1.4851 | 283.2955
|  | 1.2634 | 315.8688 | [8 empty] | 1.2986 | 314.3015 | 1.0944 | 300.5757 | 1.4855 | 283.1563
|  | 1.2637 | 315.7159 | [8 empty] | 1.299 | 314.1608 | 1.0946 | 300.2624 | 1.4859 | 283.0169
|  | 1.264 | 315.5625 | [8 empty] | 1.2993 | 314.0195 | 1.0948 | 299.9461 | 1.4864 | 282.8773
|  | 1.2643 | 315.4085 | [8 empty] | 1.2996 | 313.8778 | 1.0951 | 299.6281 | 1.4868 | 282.7375
|  | 1.2646 | 315.2538 | [8 empty] | 1.3 | 313.7357 | 1.0953 | 299.3071 | 1.4873 | 282.5976
|  | 1.265 | 315.0985 | [8 empty] | 1.3003 | 313.593 | 1.0956 | 298.9843 | 1.4877 | 282.4575
|  | 1.2653 | 314.9425 | [8 empty] | 1.3006 | 313.4497 | 1.0958 | 298.6585 | 1.4881 | 282.3173
|  | 1.2656 | 314.786 | [8 empty] | 1.301 | 313.3062 | 1.096 | 298.3296 | 1.4886 | 282.1768
|  | 1.2659 | 314.6289 | [8 empty] | 1.3013 | 313.1621 | 1.0963 | 297.9988 | 1.489 | 282.0362
|  | 1.2662 | 314.4712 | [8 empty] | 1.3016 | 313.0175 | 1.0965 | 297.6649 | 1.4894 | 281.8955
|  | 1.2665 | 314.3127 | [8 empty] | 1.302 | 312.8724 | 1.0968 | 297.3284 | 1.4899 | 281.7545
|  | 1.2669 | 314.1536 | [8 empty] | 1.3023 | 312.7268 | 1.097 | 296.9899 | 1.4903 | 281.6134
|  | 1.2672 | 313.9937 | [8 empty] | 1.3026 | 312.5807 | 1.0972 | 296.6483 | 1.4907 | 281.4721
|  | 1.2675 | 313.8334 | [8 empty] | 1.303 | 312.4342 | 1.0975 | 296.3048 | 1.4912 | 281.3307
|  | 1.2678 | 313.6725 | [8 empty] | 1.3033 | 312.2871 | 1.0977 | 295.9579 | 1.4916 | 281.1891
|  | 1.2681 | 313.5108 | [8 empty] | 1.3036 | 312.1396 | 1.098 | 295.6078 | 1.4921 | 281.0473
|  | 1.2684 | 313.3485 | [8 empty] | 1.304 | 311.9914 | 1.0982 | 295.2558 | 1.4925 | 280.9053
|  | 1.2687 | 313.1854 | [8 empty] | 1.3043 | 311.8428 | 1.0984 | 294.9004 | 1.4929 | 280.7632
|  | 1.2691 | 313.0218 | [8 empty] | 1.3046 | 311.6938 | 1.0987 | 294.5416 | 1.4934 | 280.6209
|  | 1.2694 | 312.8575 | [8 empty] | 1.305 | 311.5443 | 1.0989 | 294.1809 | 1.4938 | 280.4784
|  | 1.2697 | 312.6925 | [8 empty] | 1.3053 | 311.3942 | 1.0991 | 293.8167 | 1.4942 | 280.3358
|  | 1.27 | 312.5268 | [8 empty] | 1.3056 | 311.2436 | 1.0994 | 293.4506 | 1.4947 | 280.193
|  | 1.2703 | 312.3604 | [8 empty] | 1.306 | 311.0924 | 1.0996 | 293.081 | 1.4951 | 280.05
|  | 1.2706 | 312.1932 | [8 empty] | 1.3063 | 310.9406 | 1.0999 | 292.7079 | 1.4956 | 279.9068
|  | 1.2709 | 312.0255 | [8 empty] | 1.3066 | 310.7885 | 1.1001 | 292.333 | 1.496 | 279.7635
|  | 1.2713 | 311.8571 | [8 empty] | 1.307 | 310.6358 | 1.1003 | 291.9545 | 1.4964 | 279.62
|  | 1.2716 | 311.6879 | [8 empty] | 1.3073 | 310.4826 | 1.1006 | 291.5741 | 1.4969 | 279.4763
|  | 1.2719 | 311.5179 | [8 empty] | 1.3076 | 310.3288 | 1.1008 | 291.1908 | 1.4973 | 279.3325
|  | 1.2722 | 311.3473 | [8 empty] | 1.308 | 310.1744 | 1.1011 | 290.8041 | 1.4977 | 279.1885
|  | 1.2725 | 311.1761 | [8 empty] | 1.3083 | 310.0194 | 1.1013 | 290.4156 | 1.4982 | 279.0443
|  | 1.2728 | 311.0042 | [8 empty] | 1.3086 | 309.8641 | 1.1015 | 290.0236 | 1.4986 | 278.8999
|  | 1.2731 | 310.8314 | [8 empty] | 1.3089 | 309.7081 | 1.1018 | 289.6299 | 1.499 | 278.7554
|  | 1.2735 | 310.6579 | [8 empty] | 1.3093 | 309.5517 | 1.102 | 289.2328 | 1.4995 | 278.6107
|  | 1.2738 | 310.4835 | [8 empty] | 1.3096 | 309.3946 | 1.1023 | 288.8324 | 1.4999 | 278.4658
|  | 1.2741 | 310.3085 | [8 empty] | 1.3099 | 309.2369 | 1.1025 | 288.4305 | 1.5004 | 278.3208
|  | 1.2744 | 310.1329 | [8 empty] | 1.3103 | 309.0792 | 1.1027 | 288.0253 | 1.5008 | 278.1756
|  | 1.2747 | 309.9565 | [8 empty] | 1.3106 | 308.9208 | 1.103 | 287.6188 | 1.5012 | 278.0302
|  | 1.275 | 309.7792 | [8 empty] | 1.3109 | 308.7615 | 1.1032 | 287.2092 | 1.5017 | 277.8846
|  | 1.2754 | 309.6012 | [8 empty] | 1.3113 | 308.6018 | 1.1035 | 286.7985 | 1.5021 | 277.7389
|  | 1.2757 | 309.4223 | [8 empty] | 1.3116 | 308.4416 | 1.1037 | 286.3849 | 1.5025 | 277.593
|  | 1.276 | 309.2428 | [8 empty] | 1.3119 | 308.2807 | 1.1039 | 285.9706 | 1.503 | 277.4469
|  | 1.2763 | 309.0626 | [8 empty] | 1.3123 | 308.1194 | 1.1042 | 285.5536 | 1.5034 | 277.3007
|  | 1.2766 | 308.8816 | [8 empty] | 1.3126 | 307.9575 | 1.1044 | 285.1363 | 1.5038 | 277.1542
|  | 1.2769 | 308.6997 | [8 empty] | 1.3129 | 307.7949 | 1.1046 | 284.7187 | 1.5043 | 277.0076
|  | 1.2772 | 308.5169 | [8 empty] | 1.3133 | 307.6318 | 1.1049 | 284.2991 | 1.5047 | 276.8609
|  | 1.2776 | 308.3332 | [8 empty] | 1.3136 | 307.468 | 1.1051 | 283.8799 | 1.5052 | 276.7139
|  | 1.2779 | 308.149 | [8 empty] | 1.3139 | 307.3038 | 1.1054 | 283.4592 | 1.5056 | 276.5668
|  | 1.2782 | 307.964 | [8 empty] | 1.3143 | 307.139 | 1.1056 | 283.0396 | 1.506 | 276.4195
|  | 1.2785 | 307.7781 | [8 empty] | 1.3146 | 306.9735 | 1.1058 | 282.6212 | 1.5065 | 276.272
|  | 1.2788 | 307.5913 | [8 empty] | 1.3149 | 306.8073 | 1.1061 | 282.2024 | 1.5069 | 276.1244
|  | 1.2791 | 307.4036 | [8 empty] | 1.3153 | 306.6405 | 1.1063 | 281.7859 | 1.5073 | 275.9766
|  | 1.2794 | 307.2153 | [8 empty] | 1.3156 | 306.4732 | 1.1066 | 281.372 | 1.5078 | 275.8286
|  | 1.2798 | 307.0261 | [8 empty] | 1.3159 | 306.3054 | 1.1068 | 280.9611 | 1.5082 | 275.6804
|  | 1.2801 | 306.836 | [8 empty] | 1.3163 | 306.1369 | 1.107 | 280.5539 | 1.5086 | 275.5321
|  | 1.2804 | 306.645 | [8 empty] | 1.3166 | 305.9677 | 1.1073 | 280.1509 | 1.5091 | 275.3836
|  | 1.2807 | 306.4531 | [8 empty] | 1.3169 | 305.7978 | 1.1075 | 279.7528 | 1.5095 | 275.2349
|  | 1.281 | 306.2603 | [8 empty] | 1.3173 | 305.6275 | 1.1078 | 279.3601 | 1.51 | 275.0861
|  | 1.2813 | 306.0668 | [8 empty] | 1.3176 | 305.4565 | 1.108 | 278.9737 | 1.5104 | 274.937
|  | 1.2816 | 305.8724 | [8 empty] | 1.3179 | 305.2848 | 1.1082 | 278.5943 | 1.5108 | 274.7878
|  | 1.282 | 305.677 | [8 empty] | 1.3183 | 305.1125 | 1.1085 | 278.2225 | 1.5113 | 274.6384
|  | 1.2823 | 305.4807 | [8 empty] | 1.3186 | 304.9395 | 1.1087 | 277.8592 | 1.5117 | 274.4889
|  | 1.2826 | 305.2835 | [8 empty] | 1.3189 | 304.766 | 1.109 | 277.5052 | 1.5121 | 274.3392
|  | 1.2829 | 305.0852 | [8 empty] | 1.3193 | 304.5918 | 1.1092 | 277.1631 | 1.5126 | 274.1893
|  | 1.2832 | 304.8863 | [8 empty] | 1.3196 | 304.4169 | 1.1094 | 276.8332 | 1.513 | 274.0392
|  | 1.2835 | 304.6865 | [8 empty] | 1.3199 | 304.2413 | 1.1097 | 276.5161 | 1.5135 | 273.8889
|  | 1.2839 | 304.4855 | [8 empty] | 1.3203 | 304.065 | 1.1099 | 276.2107 | 1.5139 | 273.7385
|  | 1.2842 | 304.2836 | [8 empty] | 1.3206 | 303.8883 | 1.1101 | 275.9197 | 1.5143 | 273.5879
|  | 1.2845 | 304.0807 | [8 empty] | 1.3209 | 303.7108 | 1.1104 | 275.6429 | 1.5148 | 273.4371
|  | 1.2848 | 303.8772 | [8 empty] | 1.3213 | 303.5326 | 1.1106 | 275.381 | 1.5152 | 273.2862
|  | 1.2851 | 303.6727 | [8 empty] | 1.3216 | 303.3537 | 1.1109 | 275.134 | 1.5156 | 273.135
|  | 1.2854 | 303.4671 | [8 empty] | 1.3219 | 303.174 | 1.1111 | 274.9034 | 1.5161 | 272.9837
|  | 1.2857 | 303.2604 | [8 empty] | 1.3222 | 302.9938 | 1.1113 | 274.6114 | 1.5165 | 272.8322
|  | 1.2861 | 303.0526 | [8 empty] | 1.3226 | 302.8128 | 1.1116 | 274.3337 | 1.5169 | 272.6805
|  | 1.2864 | 302.8439 | [8 empty] | 1.3229 | 302.6312 | 1.1118 | 274.0534 | 1.5174 | 272.5287
|  | 1.2867 | 302.6346 | [8 empty] | 1.3232 | 302.4488 | 1.1121 | 273.7744 | 1.5178 | 272.3767
|  | 1.287 | 302.4241 | [8 empty] | 1.3236 | 302.2657 | 1.1123 | 273.4957 | 1.5183 | 272.2245
|  | 1.2873 | 302.2125 | [8 empty] | 1.3239 | 302.082 | 1.1125 | 273.2176 | 1.5187 | 272.0722
|  | 1.2876 | 301.9998 | [8 empty] | 1.3242 | 301.8976 | 1.1128 | 272.9401 | 1.5191 | 271.9196
|  | 1.2879 | 301.786 | [8 empty] | 1.3246 | 301.7123 | 1.113 | 272.6631 | 1.5196 | 271.7669
|  | 1.2883 | 301.571 | [8 empty] | 1.3249 | 301.5263 | 1.1133 | 272.3866 | 1.52 | 271.614
|  | 1.2886 | 301.3553 | [8 empty] | 1.3252 | 301.3396 | 1.1135 | 272.1106 | 1.5204 | 271.4609
|  | 1.2889 | 301.1386 | [8 empty] | 1.3256 | 301.1523 | 1.1137 | 271.8352 | 1.5209 | 271.3077
|  | 1.2892 | 300.9208 | [8 empty] | 1.3259 | 300.9643 | 1.114 | 271.5602 | 1.5213 | 271.1542
|  | 1.2895 | 300.7017 | [8 empty] | 1.3262 | 300.7754 | 1.1142 | 271.2858 | 1.5217 | 271.0006
|  | 1.2898 | 300.4815 | [8 empty] | 1.3266 | 300.5857 | 1.1145 | 271.012 | 1.5222 | 270.8468
|  | 1.2901 | 300.2601 | [8 empty] | 1.3269 | 300.3952 | 1.1147 | 270.7386 | 1.5226 | 270.6929
|  | 1.2905 | 300.0376 | [8 empty] | 1.3272 | 300.2041 | 1.1149 | 270.4658 | 1.5231 | 270.5387
|  | 1.2908 | 299.8144 | [8 empty] | 1.3276 | 300.0122 | 1.1152 | 270.1935 | 1.5235 | 270.3844
|  | 1.2911 | 299.59 | [8 empty] | 1.3279 | 299.8193 | 1.1154 | 269.9217 | 1.5239 | 270.23
|  | 1.2914 | 299.3643 | [8 empty] | 1.3282 | 299.6255 | 1.1156 | 269.6505 | 1.5244 | 270.0753
|  | 1.2917 | 299.1373 | [8 empty] | 1.3286 | 299.4309 | 1.1159 | 269.3797 | 1.5248 | 269.9204
|  | 1.292 | 298.9091 | [8 empty] | 1.3289 | 299.2354 | 1.1161 | 269.1095 | 1.5252 | 269.7654
|  | 1.2924 | 298.6796 | [8 empty] | 1.3292 | 299.0394 | 1.1164 | 268.8398 | 1.5257 | 269.6102
|  | 1.2927 | 298.449 | [8 empty] | 1.3296 | 298.8425 | 1.1166 | 268.5706 | 1.5261 | 269.4548
|  | 1.293 | 298.2176 | [8 empty] | 1.3299 | 298.6448 | 1.1168 | 268.3019 | 1.5265 | 269.2992
|  | 1.2933 | 297.9848 | [8 empty] | 1.3302 | 298.4462 | 1.1171 | 268.0338 | 1.527 | 269.1435
|  | 1.2936 | 297.7508 | [8 empty] | 1.3306 | 298.2467 | 1.1173 | 267.7661 | 1.5274 | 268.9876
|  | 1.2939 | 297.5154 | [8 empty] | 1.3309 | 298.0467 | 1.1176 | 267.499 | 1.5279 | 268.8314
|  | 1.2942 | 297.2786 | [8 empty] | 1.3312 | 297.8457 | 1.1178 | 267.2324 | 1.5283 | 268.6751
|  | 1.2946 | 297.0407 | [8 empty] | 1.3316 | 297.6439 | 1.118 | 266.9662 | 1.5287 | 268.5187
|  | 1.2949 | 296.8019 | [8 empty] | 1.3319 | 297.4412 | 1.1183 | 266.7006 | 1.5292 | 268.3621
|  | 1.2952 | 296.5617 | [8 empty] | 1.3322 | 297.2375 | 1.1185 | 266.4355 | 1.5296 | 268.2052
|  | 1.2955 | 296.3201 | [8 empty] | 1.3326 | 297.033 | 1.1188 | 266.1709 | 1.53 | 268.0482
|  | 1.2958 | 296.0771 | [8 empty] | 1.3329 | 296.8278 | 1.119 | 265.9068 | 1.5305 | 267.891
|  | 1.2961 | 295.8327 | [8 empty] | 1.3332 | 296.6217 | 1.1192 | 265.6432 | 1.5309 | 267.7337
|  | 1.2964 | 295.5868 | [8 empty] | 1.3336 | 296.4146 | 1.1195 | 265.3801 | 1.5314 | 267.5761
|  | 1.2968 | 295.3397 | [8 empty] | 1.3339 | 296.2067 | 1.1197 | 265.1175 | 1.5318 | 267.4184
|  | 1.2971 | 295.0911 | [8 empty] | 1.3342 | 295.9977 | 1.12 | 264.8554 | 1.5322 | 267.2605
|  | 1.2974 | 294.8415 | [8 empty] | 1.3346 | 295.7878 | 1.1202 | 264.5938 | 1.5327 | 267.1024
|  | 1.2977 | 294.5905 | [8 empty] | 1.3349 | 295.5773 | 1.1204 | 264.3327 | 1.5331 | 266.9442
|  | 1.298 | 294.3379 | [8 empty] | 1.3352 | 295.3657 | 1.1207 | 264.0721 | 1.5335 | 266.7857
|  | 1.2983 | 294.0838 | [8 empty] | 1.3355 | 295.1532 | 1.1209 | 263.812 | 1.534 | 266.6271
|  | 1.2986 | 293.8283 | [8 empty] | 1.3359 | 294.9398 | 1.1211 | 263.5524 | 1.5344 | 266.4683
|  | 1.299 | 293.5713 | [8 empty] | 1.3362 | 294.7253 | 1.1214 | 263.2933 | 1.5348 | 266.3093
|  | 1.2993 | 293.3127 | [8 empty] | 1.3365 | 294.5102 | 1.1216 | 263.0347 | 1.5353 | 266.1501
|  | 1.2996 | 293.0531 | [8 empty] | 1.3369 | 294.294 | 1.1219 | 262.7765 | 1.5357 | 265.9908
|  | 1.2999 | 292.7918 | [8 empty] | 1.3372 | 294.0768 | 1.1221 | 262.5189 | 1.5362 | 265.8312
|  | 1.3002 | 292.5289 | [8 empty] | 1.3375 | 293.8585 | 1.1223 | 262.2617 | 1.5366 | 265.6715
|  | 1.3005 | 292.2644 | [8 empty] | 1.3379 | 293.6393 | 1.1226 | 262.0051 | 1.537 | 265.5116
|  | 1.3009 | 291.9984 | [8 empty] | 1.3382 | 293.419 | 1.1228 | 261.7489 | 1.5375 | 265.3515
|  | 1.3012 | 291.7307 | [8 empty] | 1.3385 | 293.198 | 1.1231 | 261.4932 | 1.5379 | 265.1912
|  | 1.3015 | 291.4613 | [8 empty] | 1.3389 | 292.9759 | 1.1233 | 261.238 | 1.5383 | 265.0308
|  | 1.3018 | 291.1902 | [8 empty] | 1.3392 | 292.7527 | 1.1235 | 260.9832 | 1.5388 | 264.8702
|  | 1.3021 | 290.918 | [8 empty] | 1.3395 | 292.5284 | 1.1238 | 260.729 | 1.5392 | 264.7093
|  | 1.3024 | 290.644 | [8 empty] | 1.3399 | 292.3031 | 1.124 | 260.4752 | 1.5396 | 264.5483
|  | 1.3027 | 290.3682 | [8 empty] | 1.3402 | 292.0766 | 1.1243 | 260.2219 | 1.5401 | 264.3872
|  | 1.3031 | 290.0908 | [8 empty] | 1.3405 | 291.8495 | 1.1245 | 259.9691 | 1.5405 | 264.2258
|  | 1.3034 | 289.8117 | [8 empty] | 1.3409 | 291.6212 | 1.1247 | 259.7168 | 1.541 | 264.0643
|  | 1.3037 | 289.5306 | [8 empty] | 1.3412 | 291.3917 | 1.125 | 259.465 | 1.5414 | 263.9025
|  | 1.304 | 289.2477 | [8 empty] | 1.3415 | 291.161 | 1.1252 | 259.2136 | 1.5418 | 263.7406
|  | 1.3043 | 288.9629 | [8 empty] | 1.3419 | 290.9293 | 1.1255 | 258.9627 | 1.5423 | 263.5785
|  | 1.3046 | 288.6769 | [8 empty] | 1.3422 | 290.6964 | 1.1257 | 258.7123 | 1.5427 | 263.4162
|  | 1.3049 | 288.3892 | [8 empty] | 1.3425 | 290.4627 | 1.1259 | 258.4623 | 1.5431 | 263.2537
|  | 1.3053 | 288.0995 | [8 empty] | 1.3429 | 290.2278 | 1.1262 | 258.2128 | 1.5436 | 263.0911
|  | 1.3056 | 287.8078 | [8 empty] | 1.3432 | 289.9917 | 1.1264 | 257.9638 | 1.544 | 262.9282
|  | 1.3059 | 287.5142 | [8 empty] | 1.3435 | 289.7543 | 1.1266 | 257.7153 | 1.5444 | 262.7652
|  | 1.3062 | 287.2185 | [8 empty] | 1.3439 | 289.5158 | 1.1269 | 257.4672 | 1.5449 | 262.6019
|  | 1.3065 | 286.9207 | [8 empty] | 1.3442 | 289.2761 | 1.1271 | 257.2196 | 1.5453 | 262.4386
|  | 1.3068 | 286.6211 | [8 empty] | 1.3445 | 289.0355 | 1.1274 | 256.9725 | 1.5458 | 262.275
|  | 1.3071 | 286.3195 | [8 empty] | 1.3449 | 288.7937 | 1.1276 | 256.7258 | 1.5462 | 262.1112
|  | 1.3075 | 286.0156 | [8 empty] | 1.3452 | 288.5505 | 1.1278 | 256.4796 | 1.5466 | 261.9473
|  | 1.3078 | 285.7104 | [8 empty] | 1.3455 | 288.306 | 1.1281 | 256.2339 | 1.5471 | 261.7831
|  | 1.3081 | 285.4029 | [8 empty] | 1.3459 | 288.0604 | 1.1283 | 255.9886 | 1.5475 | 261.6187
|  | 1.3084 | 285.0932 | [8 empty] | 1.3462 | 287.8139 | 1.1286 | 255.7438 | 1.5479 | 261.4542
|  | 1.3087 | 284.7813 | [8 empty] | 1.3465 | 287.566 | 1.1288 | 255.4994 | 1.5484 | 261.2895
|  | 1.309 | 284.4674 | [8 empty] | 1.3469 | 287.3168 | 1.129 | 255.2555 | 1.5488 | 261.1247
|  | 1.3094 | 284.1512 | [8 empty] | 1.3472 | 287.0663 | 1.1293 | 255.012 | 1.5493 | 260.9596
|  | 1.3097 | 283.8326 | [8 empty] | 1.3475 | 286.8143 | 1.1295 | 254.769 | 1.5497 | 260.7943
|  | 1.31 | 283.5117 | [8 empty] | 1.3479 | 286.5612 | 1.1298 | 254.5265 | 1.5501 | 260.6288
|  | 1.3103 | 283.1884 | [8 empty] | 1.3482 | 286.3066 | 1.13 | 254.2844 | 1.5506 | 260.4632
|  | 1.3106 | 282.8626 | [8 empty] | 1.3485 | 286.0511 | 1.1302 | 254.0428 | 1.551 | 260.2974
|  | 1.3109 | 282.5343 | [8 empty] | 1.3488 | 285.7942 | 1.1305 | 253.8016 | 1.5514 | 260.1314
|  | 1.3112 | 282.2039 | [8 empty] | 1.3492 | 285.5358 | 1.1307 | 253.5609 | 1.5519 | 259.9652
|  | 1.3116 | 281.8718 | [8 empty] | 1.3495 | 285.2759 | 1.131 | 253.3206 | 1.5523 | 259.7988
|  | 1.3119 | 281.537 | [8 empty] | 1.3498 | 285.0145 | 1.1312 | 253.0808 | 1.5527 | 259.6322
|  | 1.3122 | 281.1997 | [8 empty] | 1.3502 | 284.7519 | 1.1314 | 252.8414 | 1.5532 | 259.4654
|  | 1.3125 | 280.8597 | [8 empty] | 1.3505 | 284.4877 | 1.1317 | 252.6025 | 1.5536 | 259.2985
|  | 1.3128 | 280.5171 | [8 empty] | 1.3508 | 284.2225 | 1.1319 | 252.364 | 1.5541 | 259.1313
|  | 1.3131 | 280.1717 | [8 empty] | 1.3512 | 283.9558 | 1.1321 | 252.126 | 1.5545 | 258.964
|  | 1.3134 | 279.8235 | [8 empty] | 1.3515 | 283.6875 | 1.1324 | 251.8884 | 1.5549 | 258.7964
|  | 1.3138 | 279.473 | [8 empty] | 1.3518 | 283.4176 | 1.1326 | 251.6512 | 1.5554 | 258.6287
|  | 1.3141 | 279.1196 | [8 empty] | 1.3522 | 283.1463 | 1.1329 | 251.4145 | 1.5558 | 258.4608
|  | 1.3144 | 278.7633 | [8 empty] | 1.3525 | 282.8734 | 1.1331 | 251.1782 | 1.5562 | 258.2927
|  | 1.3147 | 278.4041 | [8 empty] | 1.3528 | 282.5989 | 1.1333 | 250.9424 | 1.5567 | 258.1244
|  | 1.315 | 278.0419 | [8 empty] | 1.3532 | 282.3226 | 1.1336 | 250.707 | 1.5571 | 257.9559
|  | 1.3153 | 277.6767 | [8 empty] | 1.3535 | 282.0453 | 1.1338 | 250.472 | 1.5575 | 257.7872
|  | 1.3156 | 277.3089 | [8 empty] | 1.3538 | 281.7663 | 1.1341 | 250.2375 | 1.558 | 257.6184
|  | 1.316 | 276.938 | [8 empty] | 1.3542 | 281.4856 | 1.1343 | 250.0034 | 1.5584 | 257.4493
|  | 1.3163 | 276.564 | [8 empty] | 1.3545 | 281.2034 | 1.1345 | 249.7698 | 1.5589 | 257.2801
|  | 1.3166 | 276.1868 | [8 empty] | 1.3548 | 280.9194 | 1.1348 | 249.5366 | 1.5593 | 257.1106
|  | 1.3169 | 275.8063 | [8 empty] | 1.3552 | 280.6336 | 1.135 | 249.3038 | 1.5597 | 256.941
|  | 1.3172 | 275.4225 | [8 empty] | 1.3555 | 280.3461 | 1.1353 | 249.0714 | 1.5602 | 256.7712
|  | 1.3175 | 275.0354 | [8 empty] | 1.3558 | 280.0567 | 1.1355 | 248.8395 | 1.5606 | 256.6011
|  | 1.3179 | 274.645 | [8 empty] | 1.3562 | 279.7654 | 1.1357 | 248.608 | 1.561 | 256.4309
|  | 1.3182 | 274.2516 | [8 empty] | 1.3565 | 279.4729 | 1.136 | 248.3769 | 1.5615 | 256.2605
|  | 1.3185 | 273.8548 | [8 empty] | 1.3568 | 279.1788 | 1.1362 | 248.1463 | 1.5619 | 256.0899
|  | 1.3188 | 273.4544 | [8 empty] | 1.3572 | 278.8828 | 1.1365 | 247.9161 | 1.5623 | 255.9192
|  | 1.3191 | 273.0505 | [8 empty] | 1.3575 | 278.5848 | 1.1367 | 247.6863 | 1.5628 | 255.7482
|  | 1.3194 | 272.643 | [8 empty] | 1.3578 | 278.2849 | 1.1369 | 247.4569 | 1.5632 | 255.577
|  | 1.3197 | 272.2318 | [8 empty] | 1.3582 | 277.983 | 1.1372 | 247.228 | 1.5637 | 255.4056
|  | 1.3201 | 271.8169 | [8 empty] | 1.3585 | 277.6793 | 1.1374 | 246.9995 | 1.5641 | 255.234
|  | 1.3204 | 271.3989 | [8 empty] | 1.3588 | 277.3737 | 1.1376 | 246.7713 | 1.5645 | 255.0623
|  | 1.3207 | 270.9771 | [8 empty] | 1.3592 | 277.0659 | 1.1379 | 246.5437 | 1.565 | 254.8903
|  | 1.321 | 270.5515 | [8 empty] | 1.3595 | 276.756 | 1.1381 | 246.3164 | 1.5654 | 254.7181
|  | 1.3213 | 270.122 | [8 empty] | 1.3598 | 276.4448 | 1.1384 | 246.0896 | 1.5658 | 254.5458
|  | 1.3216 | 269.6887 | [8 empty] | 1.3602 | 276.1314 | 1.1386 | 245.8631 | 1.5663 | 254.3732
|  | 1.3219 | 269.2514 | [8 empty] | 1.3605 | 275.8158 | 1.1388 | 245.6371 | 1.5667 | 254.2005
|  | 1.3223 | 268.8101 | [8 empty] | 1.3608 | 275.4983 | 1.1391 | 245.4115 | 1.5672 | 254.0276
|  | 1.3226 | 268.3649 | [8 empty] | 1.3612 | 275.1786 | 1.1393 | 245.1863 | 1.5676 | 253.8544
|  | 1.3229 | 267.9164 | [8 empty] | 1.3615 | 274.8566 | 1.1396 | 244.9615 | 1.568 | 253.6811
|  | 1.3232 | 267.4638 | [8 empty] | 1.3618 | 274.5323 | 1.1398 | 244.7372 | 1.5685 | 253.5075
|  | 1.3235 | 267.0073 | [8 empty] | 1.3621 | 274.2056 | 1.14 | 244.5132 | 1.5689 | 253.3338
|  | 1.3238 | 266.5468 | [8 empty] | 1.3625 | 273.8766 | 1.1403 | 244.2897 | 1.5693 | 253.1599
|  | 1.3241 | 266.0823 | [8 empty] | 1.3628 | 273.5455 | 1.1405 | 244.0666 | 1.5698 | 252.9858
|  | 1.3245 | 265.6138 | [8 empty] | 1.3631 | 273.212 | 1.1408 | 243.8438 | 1.5702 | 252.8115
|  | 1.3248 | 265.1413 | [8 empty] | 1.3635 | 272.8769 | 1.141 | 243.6215 | 1.5706 | 252.6369
|  | 1.3251 | 264.6651 | [8 empty] | 1.3638 | 272.5393 | 1.1412 | 243.3996 | 1.5711 | 252.4622
|  | 1.3254 | 264.1858 | [8 empty] | 1.3641 | 272.1991 | 1.1415 | 243.1781 | 1.5715 | 252.2873
|  | 1.3257 | 263.7028 | [8 empty] | 1.3645 | 271.8563 | 1.1417 | 242.957 | 1.572 | 252.1122
|  | 1.326 | 263.2162 | [8 empty] | 1.3648 | 271.5109 | 1.142 | 242.7363 | 1.5724 | 251.9369
|  | 1.3264 | 262.7262 | [8 empty] | 1.3651 | 271.1632 | 1.1422 | 242.516 | 1.5728 | 251.7613
|  | 1.3267 | 262.2328 | [8 empty] | 1.3655 | 270.8129 | 1.1424 | 242.2961 | 1.5733 | 251.5856
|  | 1.327 | 261.7364 | [8 empty] | 1.3658 | 270.4598 | 1.1427 | 242.0766 | 1.5737 | 251.4097
|  | 1.3273 | 261.2372 | [8 empty] | 1.3661 | 270.104 | 1.1429 | 241.8575 | 1.5741 | 251.2336
|  | 1.3276 | 260.7354 | [8 empty] | 1.3665 | 269.7453 | 1.1431 | 241.6388 | 1.5746 | 251.0572
|  | 1.3279 | 260.2314 | [8 empty] | 1.3668 | 269.3838 | 1.1434 | 241.4204 | 1.575 | 250.8807
|  | 1.3282 | 259.7255 | [8 empty] | 1.3671 | 269.0198 | 1.1436 | 241.2025 | 1.5754 | 250.704
|  | 1.3286 | 259.2197 | [8 empty] | 1.3675 | 268.6529 | 1.1439 | 240.985 | 1.5759 | 250.5271
|  | 1.3289 | 258.7133 | [8 empty] | 1.3678 | 268.2829 | 1.1439 | 240.985 | 1.5763 | 250.3499
|  | 1.3292 | 258.2064 | [8 empty] | 1.3681 | 267.9099 | 1.1439 | 240.9469 | 1.5768 | 250.1726
|  | 1.3295 | 257.6999 | [8 empty] | 1.3685 | 267.5339 | 1.144 | 240.9093 | 1.5772 | 249.9951
|  | 1.3298 | 257.1956 | [8 empty] | 1.3688 | 267.1551 | 1.144 | 240.8716 | 1.5776 | 249.8174
|  | 1.3301 | 256.6927 | [8 empty] | 1.3691 | 266.7742 | 1.1441 | 240.8339 | 1.5781 | 249.6394
|  | 1.3304 | 256.1937 | [8 empty] | 1.3695 | 266.3901 | 1.1441 | 240.7962 | 1.5785 | 249.4612
|  | 1.3308 | 255.6979 | [8 empty] | 1.3698 | 266.0026 | 1.1442 | 240.7586 | 1.5789 | 249.2829
|  | 1.3311 | 255.208 | [8 empty] | 1.3701 | 265.6119 | 1.1442 | 240.7209 | 1.5794 | 249.1044
|  | 1.3314 | 254.7236 | [8 empty] | 1.3705 | 265.2177 | 1.1442 | 240.6832 | 1.5798 | 248.9256
|  | 1.3317 | 254.2476 | [8 empty] | 1.3708 | 264.8202 | 1.1443 | 240.6456 | 1.5802 | 248.7467
|  | 1.332 | 253.781 | [8 empty] | 1.3711 | 264.4191 | 1.1443 | 240.6079 | 1.5807 | 248.5675
|  | 1.3323 | 253.325 | [8 empty] | 1.3715 | 264.015 | 1.1444 | 240.5703 | 1.5811 | 248.3881
|  | 1.3326 | 252.8812 | [8 empty] | 1.3718 | 263.6074 | 1.1444 | 240.5327 | 1.5816 | 248.2085
|  | 1.333 | 252.4511 | [8 empty] | 1.3721 | 263.1961 | 1.1445 | 240.4951 | 1.582 | 248.0288
|  | 1.3333 | 252.0375 | [8 empty] | 1.3725 | 262.7811 | 1.1445 | 240.4575 | 1.5824 | 247.8488
|  | 1.3336 | 251.6413 | [8 empty] | 1.3728 | 262.3624 | 1.1446 | 240.4199 | 1.5829 | 247.6685
|  | 1.3339 | 251.2639 | [8 empty] | 1.3731 | 261.9398 | 1.1446 | 240.3823 | 1.5833 | 247.4881
|  | 1.3342 | 250.9079 | [8 empty] | 1.3735 | 261.5134 | 1.1447 | 240.3448 | 1.5837 | 247.3075
|  | 1.3345 | 250.5738 | [8 empty] | 1.3738 | 261.0838 | 1.1447 | 240.3072 | 1.5842 | 247.1267
|  | 1.3349 | 250.2631 | [8 empty] | 1.3741 | 260.6501 | 1.1448 | 240.2696 | 1.5846 | 246.9457
|  | 1.3352 | 249.9777 | [8 empty] | 1.3745 | 260.2125 | 1.1448 | 240.2321 | 1.5851 | 246.7644
|  | 1.3355 | 249.644 | [8 empty] | 1.3748 | 259.7708 | 1.1449 | 240.1946 | 1.5855 | 246.5829
|  | 1.3358 | 249.325 | [8 empty] | 1.3751 | 259.325 | 1.1449 | 240.157 | 1.5859 | 246.4013
|  | 1.3361 | 249.0028 | [8 empty] | 1.3754 | 258.8751 | 1.145 | 240.1195 | 1.5864 | 246.2194
|  | 1.3364 | 248.6824 | [8 empty] | 1.3758 | 258.421 | 1.145 | 240.082 | 1.5868 | 246.0373
|  | 1.3367 | 248.3625 | [8 empty] | 1.3761 | 257.9627 | 1.1451 | 240.0445 | 1.5872 | 245.8551
|  | 1.3371 | 248.0434 | [8 empty] | 1.3764 | 257.5009 | 1.1451 | 240.007 | 1.5877 | 245.6726
|  | 1.3374 | 247.7251 | [8 empty] | 1.3768 | 257.0349 | 1.1452 | 239.9696 | 1.5881 | 245.4898
|  | 1.3377 | 247.4075 | [8 empty] | 1.3771 | 256.5646 | 1.1452 | 239.9321 | 1.5885 | 245.3069
|  | 1.338 | 247.0906 | [8 empty] | 1.3774 | 256.09 | 1.1453 | 239.8946 | 1.589 | 245.1238
|  | 1.3383 | 246.7745 | [8 empty] | 1.3778 | 255.6112 | 1.1453 | 239.8572 | 1.5894 | 244.9404
|  | 1.3386 | 246.4591 | [8 empty] | 1.3781 | 255.1282 | 1.1453 | 239.8197 | 1.5899 | 244.7568
|  | 1.3389 | 246.1445 | [8 empty] | 1.3784 | 254.6411 | 1.1454 | 239.7823 | 1.5903 | 244.5731
|  | 1.3393 | 245.8306 | [8 empty] | 1.3788 | 254.1497 | 1.1454 | 239.7449 | 1.5907 | 244.3891
|  | 1.3396 | 245.5174 | [8 empty] | 1.3791 | 253.6544 | 1.1455 | 239.7075 | 1.5912 | 244.2049
|  | 1.3399 | 245.205 | [8 empty] | 1.3794 | 253.156 | 1.1455 | 239.6701 | 1.5916 | 244.0205
|  | 1.3402 | 244.8933 | [8 empty] | 1.3798 | 252.6537 | 1.1456 | 239.6327 | 1.592 | 243.8359
|  | 1.3405 | 244.5823 | [8 empty] | 1.3801 | 252.1478 | 1.1456 | 239.5953 | 1.5925 | 243.651
|  | 1.3408 | 244.272 | [8 empty] | 1.3804 | 251.6384 | 1.1457 | 239.5579 | 1.5929 | 243.4659
|  | 1.3412 | 243.9625 | [8 empty] | 1.3808 | 251.1259 | 1.1457 | 239.5206 | 1.5933 | 243.2806
|  | 1.3415 | 243.6537 | [8 empty] | 1.3811 | 250.6105 | 1.1458 | 239.4832 | 1.5938 | 243.0951
|  | 1.3418 | 243.3456 | [8 empty] | 1.3814 | 250.0926 | 1.1458 | 239.4458 | 1.5942 | 242.9094
|  | 1.3421 | 243.0382 | [8 empty] | 1.3818 | 249.5725 | 1.1459 | 239.4085 | 1.5947 | 242.7235
|  | 1.3424 | 242.7315 | [8 empty] | 1.3821 | 249.0509 | 1.1459 | 239.3712 | 1.5951 | 242.5374
|  | 1.3427 | 242.4256 | [8 empty] | 1.3824 | 248.5296 | 1.146 | 239.3339 | 1.5955 | 242.351
|  | 1.343 | 242.1203 | [8 empty] | 1.3828 | 248.0076 | 1.146 | 239.2965 | 1.596 | 242.1644
|  | 1.3434 | 241.8158 | [8 empty] | 1.3831 | 247.4868 | 1.1461 | 239.2592 | 1.5964 | 241.9776
|  | 1.3437 | 241.5119 | [8 empty] | 1.3834 | 246.9685 | 1.1461 | 239.222 | 1.5968 | 241.7906
|  | 1.344 | 241.2088 | [8 empty] | 1.3838 | 246.452 | 1.1462 | 239.1847 | 1.5973 | 241.6034
|  | 1.3443 | 240.9064 | [8 empty] | 1.3841 | 245.9402 | 1.1462 | 239.1474 | 1.5977 | 241.4159
|  | 1.3446 | 240.6046 | [8 empty] | 1.3844 | 245.4323 | 1.1463 | 239.1101 | 1.5981 | 241.2282
|  | 1.3449 | 240.3036 | [8 empty] | 1.3848 | 244.9318 | 1.1463 | 239.0729 | 1.5986 | 241.0403
|  | 1.3452 | 240.0033 | [8 empty] | 1.3851 | 244.4395 | 1.1464 | 239.0356 | 1.599 | 240.8522
|  | 1.3456 | 239.7036 | [8 empty] | 1.3854 | 243.9569 | 1.1464 | 238.9984 | 1.5995 | 240.6638
|  | 1.3459 | 239.4047 | [8 empty] | 1.3858 | 243.4859 | 1.1464 | 238.9612 | 1.5999 | 240.4752
|  | 1.3462 | 239.1064 | [8 empty] | 1.3861 | 243.0281 | 1.1465 | 238.9239 | 1.6003 | 240.2864
|  | 1.3465 | 238.8088 | [8 empty] | 1.3864 | 242.5869 | 1.1465 | 238.8867 | 1.6008 | 240.0974
|  | 1.3468 | 238.5119 | [8 empty] | 1.3868 | 242.1635 | 1.1466 | 238.8495 | 1.6012 | 239.9082
|  | 1.3471 | 238.2157 | [8 empty] | 1.3871 | 241.7592 | 1.1466 | 238.8123 | 1.6016 | 239.7188
|  | 1.3474 | 237.9201 | [8 empty] | 1.3874 | 241.3766 | 1.1467 | 238.7752 | 1.6021 | 239.529
|  | 1.3478 | 237.6253 | [8 empty] | 1.3878 | 241.0184 | 1.1467 | 238.738 | 1.6025 | 239.3391
|  | 1.3481 | 237.3311 | [8 empty] | 1.3881 | 240.6856 | 1.1468 | 238.7008 | 1.603 | 239.1489
|  | 1.3484 | 237.0376 | [8 empty] | 1.3884 | 240.3805 | 1.1468 | 238.6637 | 1.6034 | 238.9586
|  | 1.3487 | 236.7448 | [8 empty] | 1.3887 | 240.1034 | 1.1469 | 238.6265 | 1.6038 | 238.768
|  | 1.349 | 236.4526 | [8 empty] | 1.3891 | 239.7711 | 1.1469 | 238.5894 | 1.6043 | 238.5772
|  | 1.3493 | 236.1611 | [8 empty] | 1.3894 | 239.4578 | 1.147 | 238.5523 | 1.6047 | 238.3861
|  | 1.3497 | 235.8703 | [8 empty] | 1.3897 | 239.1392 | 1.147 | 238.5151 | 1.6051 | 238.1948
|  | 1.35 | 235.5801 | [8 empty] | 1.3901 | 238.8234 | 1.1471 | 238.478 | 1.6056 | 238.0033
|  | 1.3503 | 235.2906 | [8 empty] | 1.3904 | 238.5077 | 1.1471 | 238.4409 | 1.606 | 237.8116
|  | 1.3506 | 235.0017 | [8 empty] | 1.3907 | 238.193 | 1.1472 | 238.4039 | 1.6064 | 237.6196
|  | 1.3509 | 234.7135 | [8 empty] | 1.3911 | 237.879 | 1.1472 | 238.3668 | 1.6069 | 237.4274
|  | 1.3512 | 234.426 | [8 empty] | 1.3914 | 237.5657 | 1.1473 | 238.3297 | 1.6073 | 237.235
|  | 1.3515 | 234.1391 | [8 empty] | 1.3917 | 237.2532 | 1.1473 | 238.2926 | 1.6078 | 237.0423
|  | 1.3519 | 233.8529 | [8 empty] | 1.3921 | 236.9414 | 1.1474 | 238.2556 | 1.6082 | 236.8494
|  | 1.3522 | 233.5673 | [8 empty] | 1.3924 | 236.6304 | 1.1474 | 238.2185 | 1.6086 | 236.6563
|  | 1.3525 | 233.2824 | [8 empty] | 1.3927 | 236.3202 | 1.1475 | 238.1815 | 1.6091 | 236.4629
|  | 1.3528 | 232.9981 | [8 empty] | 1.3931 | 236.0106 | 1.1475 | 238.1445 | 1.6095 | 236.2693
|  | 1.3531 | 232.7144 | [8 empty] | 1.3934 | 235.7018 | 1.1475 | 238.1075 | 1.6099 | 236.0754
|  | 1.3534 | 232.4314 | [8 empty] | 1.3937 | 235.3938 | 1.1476 | 238.0705 | 1.6104 | 235.8814
|  | 1.3537 | 232.149 | [8 empty] | 1.3941 | 235.0865 | 1.1476 | 238.0335 | 1.6108 | 235.6871
|  | 1.3541 | 231.8673 | [8 empty] | 1.3944 | 234.7799 | 1.1477 | 237.9965 | 1.6112 | 235.4925
|  | 1.3544 | 231.5862 | [8 empty] | 1.3947 | 234.474 | 1.1477 | 237.9595 | 1.6117 | 235.2978
|  | 1.3547 | 231.3058 | [8 empty] | 1.3951 | 234.1689 | 1.1478 | 237.9225 | 1.6121 | 235.1027
|  | 1.355 | 231.0259 | [8 empty] | 1.3954 | 233.8645 | 1.1478 | 237.8855 | 1.6126 | 234.9074
|  | 1.3553 | 230.7467 | [8 empty] | 1.3957 | 233.5608 | 1.1479 | 237.8486 | 1.613 | 234.712
|  | 1.3556 | 230.4682 | [8 empty] | 1.3961 | 233.2579 | 1.1479 | 237.8117 | 1.6134 | 234.5162
|  | 1.3559 | 230.1902 | [8 empty] | 1.3964 | 232.9556 | 1.148 | 237.7747 | 1.6139 | 234.3203
|  | 1.3563 | 229.9129 | [8 empty] | 1.3967 | 232.6541 | 1.148 | 237.7378 | 1.6143 | 234.124
|  | 1.3566 | 229.6362 | [8 empty] | 1.3971 | 232.3533 | 1.1481 | 237.7009 | 1.6147 | 233.9276
|  | 1.3569 | 229.3601 | [8 empty] | 1.3974 | 232.0532 | 1.1481 | 237.664 | 1.6152 | 233.7309
|  | 1.3572 | 229.0847 | [8 empty] | 1.3977 | 231.7538 | 1.1482 | 237.6271 | 1.6156 | 233.534
|  | 1.3575 | 228.8098 | [8 empty] | 1.3981 | 231.4551 | 1.1482 | 237.5902 | 1.616 | 233.3368
|  | 1.3578 | 228.5356 | [8 empty] | 1.3984 | 231.1572 | 1.1483 | 237.5533 | 1.6165 | 233.1394
|  | 1.3582 | 228.262 | [8 empty] | 1.3987 | 230.8599 | 1.1483 | 237.5164 | 1.6169 | 232.9416
|  | 1.3585 | 227.9889 | [8 empty] | 1.3991 | 230.5633 | 1.1484 | 237.4796 | 1.6174 | 232.7437
|  | 1.3588 | 227.7165 | [8 empty] | 1.3994 | 230.2675 | 1.1484 | 237.4427 | 1.6178 | 232.5456
|  | 1.3591 | 227.4448 | [8 empty] | 1.3997 | 229.9723 | 1.1485 | 237.4059 | 1.6182 | 232.3472
|  | 1.3594 | 227.1736 | [8 empty] | 1.4001 | 229.6778 | 1.1485 | 237.369 | 1.6187 | 232.1486
|  | 1.3597 | 226.903 | [8 empty] | 1.4004 | 229.384 | 1.1486 | 237.3322 | 1.6191 | 231.9496
|  | 1.36 | 226.633 | [8 empty] | 1.4007 | 229.0909 | 1.1486 | 237.2954 | 1.6195 | 231.7505
|  | 1.3604 | 226.3636 | [8 empty] | 1.4011 | 228.7985 | 1.1486 | 237.2586 | 1.62 | 231.5511
|  | 1.3607 | 226.0948 | [8 empty] | 1.4014 | 228.5068 | 1.1487 | 237.2218 | 1.6204 | 231.3515
|  | 1.361 | 225.8266 | [8 empty] | 1.4017 | 228.2158 | 1.1487 | 237.185 | 1.6209 | 231.1516
|  | 1.3613 | 225.559 | [8 empty] | 1.402 | 227.9254 | 1.1488 | 237.1482 | 1.6213 | 230.9514
|  | 1.3616 | 225.292 | [8 empty] | 1.4024 | 227.6357 | 1.1488 | 237.1114 | 1.6217 | 230.751
|  | 1.3619 | 225.0256 | [8 empty] | 1.4027 | 227.3467 | 1.1489 | 237.0747 | 1.6222 | 230.5504
|  | 1.3622 | 224.7598 | [8 empty] | 1.403 | 227.0584 | 1.1489 | 237.0379 | 1.6226 | 230.3495
|  | 1.3626 | 224.4945 | [8 empty] | 1.4034 | 226.7707 | 1.149 | 237.0012 | 1.623 | 230.1483
|  | 1.3629 | 224.2299 | [8 empty] | 1.4037 | 226.4837 | 1.149 | 236.9645 | 1.6235 | 229.9468
|  | 1.3632 | 223.9658 | [8 empty] | 1.404 | 226.1974 | 1.1491 | 236.9277 | 1.6239 | 229.7451
|  | 1.3635 | 223.7023 | [8 empty] | 1.4044 | 225.9118 | 1.1491 | 236.891 | 1.6243 | 229.5432
|  | 1.3638 | 223.4394 | [8 empty] | 1.4047 | 225.6268 | 1.1492 | 236.8543 | 1.6248 | 229.3411
|  | 1.3641 | 223.177 | [8 empty] | 1.405 | 225.3424 | 1.1492 | 236.8176 | 1.6252 | 229.1386
|  | 1.3644 | 222.9153 | [8 empty] | 1.4054 | 225.0588 | 1.1493 | 236.7809 | 1.6257 | 228.9359
|  | 1.3648 | 222.6541 | [8 empty] | 1.4057 | 224.7757 | 1.1493 | 236.7442 | 1.6261 | 228.7329
|  | 1.3651 | 222.3934 | [8 empty] | 1.406 | 224.4934 | 1.1494 | 236.7076 | 1.6265 | 228.5297
|  | 1.3654 | 222.1334 | [8 empty] | 1.4064 | 224.2117 | 1.1494 | 236.6709 | 1.627 | 228.3263
|  | 1.3657 | 221.8739 | [8 empty] | 1.4067 | 223.9306 | 1.1495 | 236.6342 | 1.6274 | 228.1225
|  | 1.366 | 221.615 | [8 empty] | 1.407 | 223.6502 | 1.1495 | 236.5976 | 1.6278 | 227.9185
|  | 1.3663 | 221.3566 | [8 empty] | 1.4074 | 223.3704 | 1.1496 | 236.561 | 1.6283 | 227.7142
|  | 1.3667 | 221.0988 | [8 empty] | 1.4077 | 223.0913 | 1.1496 | 236.5243 | 1.6287 | 227.5097
|  | 1.367 | 220.8416 | [8 empty] | 1.408 | 222.8129 | 1.1497 | 236.4877 | 1.6291 | 227.3049
|  | 1.3673 | 220.5849 | [8 empty] | 1.4084 | 222.535 | 1.1497 | 236.4511 | 1.6296 | 227.0998
|  | 1.3676 | 220.3288 | [8 empty] | 1.4087 | 222.2578 | 1.1497 | 236.4145 | 1.63 | 226.8945
|  | 1.3679 | 220.0733 | [8 empty] | 1.409 | 221.9813 | 1.1498 | 236.3779 | 1.6305 | 226.6889
|  | 1.3682 | 219.8183 | [8 empty] | 1.4094 | 221.7053 | 1.1498 | 236.3413 | 1.6309 | 226.483
|  | 1.3685 | 219.5638 | [8 empty] | 1.4097 | 221.43 | 1.1499 | 236.3048 | 1.6313 | 226.2769
|  | 1.3689 | 219.3099 | [8 empty] | 1.41 | 221.1554 | 1.1499 | 236.2682 | 1.6318 | 226.0705
|  | 1.3692 | 219.0566 | [8 empty] | 1.4104 | 220.8813 | 1.15 | 236.2316 | 1.6322 | 225.8638
|  | 1.3695 | 218.8038 | [8 empty] | 1.4107 | 220.6079 | 1.15 | 236.1951 | 1.6326 | 225.6568
|  | 1.3698 | 218.5515 | [8 empty] | 1.411 | 220.3351 | 1.1501 | 236.1586 | 1.6331 | 225.4495
|  | 1.3701 | 218.2998 | [8 empty] | 1.4114 | 220.063 | 1.1501 | 236.122 | 1.6335 | 225.2421
|  | 1.3704 | 218.0487 | [8 empty] | 1.4117 | 219.7914 | 1.1502 | 236.0855 | 1.6339 | 225.0343
|  | 1.3707 | 217.798 | [8 empty] | 1.412 | 219.5205 | 1.1502 | 236.049 | 1.6344 | 224.8262
|  | 1.3711 | 217.5479 | [8 empty] | 1.4124 | 219.2502 | 1.1503 | 236.0125 | 1.6348 | 224.6178
|  | 1.3711 | 217.5479 | [8 empty] | 1.4127 | 218.9805 | 1.1503 | 235.976 | 1.6353 | 224.4092
|  | 1.3711 | 217.4974 | [8 empty] | 1.413 | 218.7114 | 1.1504 | 235.9395 | 1.6357 | 224.2003
|  | 1.3712 | 217.4468 | [8 empty] | 1.4134 | 218.4429 | 1.1504 | 235.9031 | 1.6361 | 223.9911
|  | 1.3712 | 217.3961 | [8 empty] | 1.4137 | 218.1751 | 1.1505 | 235.8666 | 1.6366 | 223.7817
|  | 1.3713 | 217.3455 | [8 empty] | 1.414 | 217.9078 | 1.1505 | 235.8301 | 1.637 | 223.5719
|  | 1.3714 | 217.2949 | [8 empty] | 1.4144 | 217.6412 | 1.1506 | 235.7937 | 1.6374 | 223.3618
|  | 1.3714 | 217.2443 | [8 empty] | 1.4147 | 217.3751 | 1.1506 | 235.7573 | 1.6379 | 223.1515
|  | 1.3715 | 217.1937 | [8 empty] | 1.415 | 217.1096 | 1.1507 | 235.7208 | 1.6383 | 222.941
|  | 1.3716 | 217.143 | [8 empty] | 1.4153 | 216.8448 | 1.1507 | 235.6844 | 1.6388 | 222.7301
|  | 1.3716 | 217.0924 | [8 empty] | 1.4157 | 216.5805 | 1.1508 | 235.648 | 1.6392 | 222.5189
|  | 1.3717 | 217.0418 | [8 empty] | 1.416 | 216.3169 | 1.1508 | 235.6116 | 1.6396 | 222.3074
|  | 1.3718 | 216.9912 | [8 empty] | 1.4163 | 216.0538 | 1.1508 | 235.5752 | 1.6401 | 222.0957
|  | 1.3718 | 216.9406 | [8 empty] | 1.4167 | 215.7913 | 1.1509 | 235.5388 | 1.6405 | 221.8836
|  | 1.3719 | 216.8899 | [8 empty] | 1.417 | 215.5294 | 1.1509 | 235.5024 | 1.6409 | 221.6713
|  | 1.3719 | 216.8393 | [8 empty] | 1.4173 | 215.2681 | 1.151 | 235.4661 | 1.6414 | 221.4586
|  | 1.372 | 216.7887 | [8 empty] | 1.4177 | 215.0074 | 1.151 | 235.4297 | 1.6418 | 221.2457
|  | 1.3721 | 216.738 | [8 empty] | 1.418 | 214.7472 | 1.1511 | 235.3934 | 1.6422 | 221.0325
|  | 1.3721 | 216.6874 | [8 empty] | 1.4183 | 214.4877 | 1.1511 | 235.357 | 1.6427 | 220.819
|  | 1.3722 | 216.6368 | [8 empty] | 1.4187 | 214.2287 | 1.1512 | 235.3207 | 1.6431 | 220.6052
|  | 1.3723 | 216.5862 | [8 empty] | 1.419 | 213.9703 | 1.1512 | 235.2844 | 1.6436 | 220.391
|  | 1.3723 | 216.5355 | [8 empty] | 1.4193 | 213.7125 | 1.1513 | 235.2481 | 1.644 | 220.1766
|  | 1.3724 | 216.4849 | [8 empty] | 1.4197 | 213.4552 | 1.1513 | 235.2118 | 1.6444 | 219.9619
|  | 1.3724 | 216.4342 | [8 empty] | 1.42 | 213.1985 | 1.1514 | 235.1755 | 1.6449 | 219.7469
|  | 1.3725 | 216.3836 | [8 empty] | 1.4203 | 212.9424 | 1.1514 | 235.1392 | 1.6453 | 219.5316
|  | 1.3726 | 216.333 | [8 empty] | 1.4207 | 212.6869 | 1.1515 | 235.1029 | 1.6457 | 219.3159
|  | 1.3726 | 216.2823 | [8 empty] | 1.421 | 212.4319 | 1.1515 | 235.0666 | 1.6462 | 219.0999
|  | 1.3727 | 216.2317 | [8 empty] | 1.4213 | 212.1775 | 1.1516 | 235.0304 | 1.6466 | 218.8837
|  | 1.3728 | 216.181 | [8 empty] | 1.4217 | 211.9236 | 1.1516 | 234.9941 | 1.647 | 218.6672
|  | 1.3728 | 216.1304 | [8 empty] | 1.422 | 211.6703 | 1.1517 | 234.9579 | 1.6475 | 218.4503
|  | 1.3729 | 216.0798 | [8 empty] | 1.4223 | 211.4176 | 1.1517 | 234.9216 | 1.6479 | 218.2331
|  | 1.3729 | 216.0291 | [8 empty] | 1.4227 | 211.1654 | 1.1518 | 234.8854 | 1.6484 | 218.0156
|  | 1.373 | 215.9785 | [8 empty] | 1.423 | 210.9138 | 1.1518 | 234.8492 | 1.6488 | 217.7978
|  | 1.3731 | 215.9278 | [8 empty] | 1.4233 | 210.6627 | 1.1519 | 234.813 | 1.6492 | 217.5797
|  | 1.3731 | 215.8772 | [8 empty] | 1.4237 | 210.4122 | 1.1519 | 234.7768 | 1.6497 | 217.3613
|  | 1.3732 | 215.8265 | [8 empty] | 1.424 | 210.1623 | 1.1519 | 234.7406 | 1.6501 | 217.1425
|  | 1.3733 | 215.7759 | [8 empty] | 1.4243 | 209.9128 | 1.152 | 234.7044 | 1.6505 | 216.9234
|  | 1.3733 | 215.7252 | [8 empty] | 1.4247 | 209.664 | 1.152 | 234.6682 | 1.651 | 216.7041
|  | 1.3734 | 215.6745 | [8 empty] | 1.425 | 209.4156 | 1.1521 | 234.6321 | 1.6514 | 216.4844
|  | 1.3735 | 215.6239 | [8 empty] | 1.4253 | 209.1679 | 1.1521 | 234.5959 | 1.6518 | 216.2643
|  | 1.3735 | 215.5732 | [8 empty] | 1.4257 | 208.9206 | 1.1522 | 234.5598 | 1.6523 | 216.0439
|  | 1.3736 | 215.5226 | [8 empty] | 1.426 | 208.6739 | 1.1522 | 234.5237 | 1.6527 | 215.8231
|  | 1.3736 | 215.4719 | [8 empty] | 1.4263 | 208.4277 | 1.1523 | 234.4875 | 1.6532 | 215.602
|  | 1.3737 | 215.4212 | [8 empty] | 1.4263 | 208.4277 | 1.1523 | 234.4514 | 1.6536 | 215.3807
|  | 1.3738 | 215.3706 | [8 empty] | 1.4264 | 208.3728 | 1.1524 | 234.4153 | 1.654 | 215.159
|  | 1.3738 | 215.3199 | [8 empty] | 1.4265 | 208.3173 | 1.1524 | 234.3792 | 1.6545 | 214.9369
|  | 1.3739 | 215.2692 | [8 empty] | 1.4265 | 208.2619 | 1.1525 | 234.3431 | 1.6549 | 214.7145
|  | 1.374 | 215.2186 | [8 empty] | 1.4266 | 208.2065 | 1.1525 | 234.307 | 1.6553 | 214.4917
|  | 1.374 | 215.1679 | [8 empty] | 1.4267 | 208.151 | 1.1526 | 234.2709 | 1.6558 | 214.2687
|  | 1.3741 | 215.1172 | [8 empty] | 1.4267 | 208.0956 | 1.1526 | 234.2349 | 1.6562 | 214.0453
|  | 1.3741 | 215.0665 | [8 empty] | 1.4268 | 208.0401 | 1.1527 | 234.1988 | 1.6567 | 213.8215
|  | 1.3742 | 215.0159 | [8 empty] | 1.4269 | 207.9847 | 1.1527 | 234.1628 | 1.6571 | 213.5974
|  | 1.3743 | 214.9652 | [8 empty] | 1.4269 | 207.9292 | 1.1528 | 234.1267 | 1.6575 | 213.3729
|  | 1.3743 | 214.9145 | [8 empty] | 1.427 | 207.8737 | 1.1528 | 234.0907 | 1.658 | 213.1481
|  | 1.3744 | 214.8638 | [8 empty] | 1.4271 | 207.8182 | 1.1529 | 234.0547 | 1.6584 | 212.923
|  | 1.3745 | 214.8131 | [8 empty] | 1.4271 | 207.7627 | 1.1529 | 234.0187 | 1.6588 | 212.6975
|  | 1.3745 | 214.7624 | [8 empty] | 1.4272 | 207.7072 | 1.153 | 233.9827 | 1.6593 | 212.4716
|  | 1.3746 | 214.7117 | [8 empty] | 1.4273 | 207.6517 | 1.153 | 233.9467 | 1.6597 | 212.2453
|  | 1.3746 | 214.661 | [8 empty] | 1.4273 | 207.5962 | 1.153 | 233.9107 | 1.6601 | 212.0188
|  | 1.3747 | 214.6103 | [8 empty] | 1.4274 | 207.5406 | 1.1531 | 233.8747 | 1.6606 | 211.7919
|  | 1.3748 | 214.5597 | [8 empty] | 1.4275 | 207.4851 | 1.1531 | 233.8387 | 1.661 | 211.5646
|  | 1.3748 | 214.509 | [8 empty] | 1.4275 | 207.4295 | 1.1532 | 233.8028 | 1.6615 | 211.337
|  | 1.3749 | 214.4582 | [8 empty] | 1.4276 | 207.374 | 1.1532 | 233.7668 | 1.6619 | 211.1089
|  | 1.375 | 214.4075 | [8 empty] | 1.4277 | 207.3184 | 1.1533 | 233.7309 | 1.6623 | 210.8805
|  | 1.375 | 214.3568 | [8 empty] | 1.4277 | 207.2628 | 1.1533 | 233.695 | 1.6628 | 210.6518
|  | 1.3751 | 214.3061 | [8 empty] | 1.4278 | 207.2073 | 1.1534 | 233.659 | 1.6632 | 210.4227
|  | 1.3752 | 214.2554 | [8 empty] | 1.4278 | 207.1517 | 1.1534 | 233.6231 | 1.6636 | 210.1931
|  | 1.3752 | 214.2047 | [8 empty] | 1.4279 | 207.0961 | 1.1535 | 233.5872 | 1.6641 | 209.9632
|  | 1.3753 | 214.154 | [8 empty] | 1.428 | 207.0404 | 1.1535 | 233.5513 | 1.6645 | 209.733
|  | 1.3753 | 214.1033 | [8 empty] | 1.428 | 206.9848 | 1.1536 | 233.5154 | 1.6649 | 209.5024
|  | 1.3754 | 214.0526 | [8 empty] | 1.4281 | 206.9292 | 1.1536 | 233.4795 | 1.6654 | 209.2714
|  | 1.3755 | 214.0018 | [8 empty] | 1.4282 | 206.8735 | 1.1537 | 233.4437 | 1.6658 | 209.0399
|  | 1.3755 | 213.9511 | [8 empty] | 1.4282 | 206.8179 | 1.1537 | 233.4078 | 1.6663 | 208.8081
|  | 1.3756 | 213.9004 | [8 empty] | 1.4283 | 206.7622 | 1.1538 | 233.3719 | 1.6667 | 208.5759
|  | 1.3757 | 213.8497 | [8 empty] | 1.4284 | 206.7066 | 1.1538 | 233.3361 | 1.6671 | 208.3434
|  | 1.3757 | 213.7989 | [8 empty] | 1.4284 | 206.6509 | 1.1539 | 233.3003 | 1.6676 | 208.1104
|  | 1.3758 | 213.7482 | [8 empty] | 1.4285 | 206.5952 | 1.1539 | 233.2644 | 1.668 | 207.877
|  | 1.3758 | 213.6974 | [8 empty] | 1.4286 | 206.5395 | 1.154 | 233.2286 | 1.6684 | 207.6432
|  | 1.3759 | 213.6467 | [8 empty] | 1.4286 | 206.4838 | 1.154 | 233.1928 | 1.6689 | 207.409
|  | 1.376 | 213.596 | [8 empty] | 1.4287 | 206.4281 | 1.1541 | 233.157 | 1.6693 | 207.1745
|  | 1.376 | 213.5452 | [8 empty] | 1.4288 | 206.3724 | 1.1541 | 233.1212 | 1.6698 | 206.9395
|  | 1.3761 | 213.4945 | [8 empty] | 1.4288 | 206.3166 | 1.1541 | 233.0854 | 1.6702 | 206.7041
|  | 1.3762 | 213.4437 | [8 empty] | 1.4289 | 206.2609 | 1.1542 | 233.0496 | 1.6706 | 206.4682
|  | 1.3762 | 213.393 | [8 empty] | 1.429 | 206.2052 | 1.1542 | 233.0139 | 1.6711 | 206.2321
|  | 1.3763 | 213.3422 | [8 empty] | 1.429 | 206.1494 | 1.1543 | 232.9781 | 1.6715 | 205.9955
|  | 1.3763 | 213.2915 | [8 empty] | 1.4291 | 206.0936 | 1.1543 | 232.9424 | 1.6719 | 205.7584
|  | 1.3764 | 213.2407 | [8 empty] | 1.4292 | 206.0379 | 1.1544 | 232.9066 | 1.6724 | 205.5209
|  | 1.3765 | 213.1899 | [8 empty] | 1.4292 | 205.9821 | 1.1544 | 232.8709 | 1.6728 | 205.2829
|  | 1.3765 | 213.1392 | [8 empty] | 1.4293 | 205.9263 | 1.1545 | 232.8352 | 1.6732 | 205.0445
|  | 1.3766 | 213.0884 | [8 empty] | 1.4294 | 205.8705 | 1.1545 | 232.7995 | 1.6737 | 204.8058
|  | 1.3767 | 213.0376 | [8 empty] | 1.4294 | 205.8147 | 1.1546 | 232.7637 | 1.6741 | 204.5666
|  | 1.3767 | 212.9869 | [8 empty] | 1.4295 | 205.7589 | 1.1546 | 232.728 | 1.6746 | 204.3269
|  | 1.3768 | 212.9361 | [8 empty] | 1.4296 | 205.703 | 1.1547 | 232.6924 | 1.675 | 204.0868
|  | 1.3769 | 212.8853 | [8 empty] | 1.4296 | 205.6472 | 1.1547 | 232.6567 | 1.6754 | 203.8462
|  | 1.3769 | 212.8345 | [8 empty] | 1.4297 | 205.5913 | 1.1548 | 232.621 | 1.6759 | 203.6053
|  | 1.377 | 212.7837 | [8 empty] | 1.4298 | 205.5355 | 1.1548 | 232.5853 | 1.6763 | 203.3639
|  | 1.377 | 212.7329 | [8 empty] | 1.4298 | 205.4796 | 1.1549 | 232.5497 | 1.6767 | 203.122
|  | 1.3771 | 212.6821 | [8 empty] | 1.4299 | 205.4238 | 1.1549 | 232.514 | 1.6772 | 202.8795
|  | 1.3772 | 212.6314 | [8 empty] | 1.43 | 205.3679 | 1.155 | 232.4784 | 1.6776 | 202.6367
|  | 1.3772 | 212.5806 | [8 empty] | 1.43 | 205.312 | 1.155 | 232.4428 | 1.678 | 202.3935
|  | 1.3773 | 212.5298 | [8 empty] | 1.4301 | 205.2561 | 1.1551 | 232.4072 | 1.6785 | 202.1498
|  | 1.3774 | 212.4789 | [8 empty] | 1.4302 | 205.2002 | 1.1551 | 232.3715 | 1.6789 | 201.9056
|  | 1.3774 | 212.4281 | [8 empty] | 1.4302 | 205.1443 | 1.1552 | 232.3359 | 1.6794 | 201.6609
|  | 1.3775 | 212.3773 | [8 empty] | 1.4303 | 205.0884 | 1.1552 | 232.3003 | 1.6798 | 201.4156
|  | 1.3775 | 212.3265 | [8 empty] | 1.4304 | 205.0324 | 1.1552 | 232.2648 | 1.6802 | 201.17
|  | 1.3776 | 212.2757 | [8 empty] | 1.4304 | 204.9765 | 1.1553 | 232.2292 | 1.6807 | 200.924
|  | 1.3777 | 212.2249 | [8 empty] | 1.4305 | 204.9205 | 1.1553 | 232.1936 | 1.6811 | 200.6774
|  | 1.3777 | 212.1741 | [8 empty] | 1.4306 | 204.8646 | 1.1554 | 232.1581 | 1.6815 | 200.4303
|  | 1.3778 | 212.1232 | [8 empty] | 1.4306 | 204.8086 | 1.1554 | 232.1225 | 1.682 | 200.1826
|  | 1.3779 | 212.0724 | [8 empty] | 1.4307 | 204.7526 | 1.1555 | 232.087 | 1.6824 | 199.9344
|  | 1.3779 | 212.0216 | [8 empty] | 1.4308 | 204.6966 | 1.1555 | 232.0514 | 1.6828 | 199.6858
|  | 1.378 | 211.9707 | [8 empty] | 1.4308 | 204.6407 | 1.1556 | 232.0159 | 1.6833 | 199.4367
|  | 1.378 | 211.9199 | [8 empty] | 1.4309 | 204.5847 | 1.1556 | 231.9804 | 1.6837 | 199.1871
|  | 1.3781 | 211.869 | [8 empty] | 1.431 | 204.5287 | 1.1557 | 231.9449 | 1.6842 | 198.9369
|  | 1.3782 | 211.8182 | [8 empty] | 1.431 | 204.4726 | 1.1557 | 231.9094 | 1.6846 | 198.6861
|  | 1.3782 | 211.7673 | [8 empty] | 1.4311 | 204.4166 | 1.1558 | 231.8739 | 1.685 | 198.435
|  | 1.3783 | 211.7165 | [8 empty] | 1.4312 | 204.3606 | 1.1558 | 231.8384 | 1.6855 | 198.1833
|  | 1.3784 | 211.6656 | [8 empty] | 1.4312 | 204.3045 | 1.1559 | 231.8029 | 1.6859 | 197.931
|  | 1.3784 | 211.6148 | [8 empty] | 1.4313 | 204.2485 | 1.1559 | 231.7675 | 1.6863 | 197.6782
|  | 1.3785 | 211.5639 | [8 empty] | 1.4314 | 204.1924 | 1.156 | 231.732 | 1.6868 | 197.4248
|  | 1.3786 | 211.513 | [8 empty] | 1.4314 | 204.1364 | 1.156 | 231.6966 | 1.6872 | 197.171
|  | 1.3786 | 211.4622 | [8 empty] | 1.4315 | 204.0803 | 1.1561 | 231.6611 | 1.6877 | 196.9166
|  | 1.3787 | 211.4113 | [8 empty] | 1.4316 | 204.0242 | 1.1561 | 231.6257 | 1.6881 | 196.6616
|  | 1.3787 | 211.3604 | [8 empty] | 1.4316 | 203.9681 | 1.1562 | 231.5903 | 1.6885 | 196.406
|  | 1.3788 | 211.3095 | [8 empty] | 1.4317 | 203.912 | 1.1562 | 231.5548 | 1.689 | 196.1498
|  | 1.3789 | 211.2586 | [8 empty] | 1.4318 | 203.8559 | 1.1563 | 231.5194 | 1.6894 | 195.893
|  | 1.3789 | 211.2077 | [8 empty] | 1.4318 | 203.7998 | 1.1563 | 231.484 | 1.6898 | 195.6358
|  | 1.379 | 211.1569 | [8 empty] | 1.4319 | 203.7437 | 1.1563 | 231.4487 | 1.6903 | 195.378
|  | 1.3791 | 211.106 | [8 empty] | 1.432 | 203.6875 | 1.1564 | 231.4133 | 1.6907 | 195.1195
|  | 1.3791 | 211.055 | [8 empty] | 1.432 | 203.6314 | 1.1564 | 231.3779 | 1.6911 | 194.8603
|  | 1.3792 | 211.0041 | [8 empty] | 1.4321 | 203.5753 | 1.1565 | 231.3425 | 1.6916 | 194.6006
|  | 1.3792 | 210.9532 | [8 empty] | 1.4322 | 203.5191 | 1.1565 | 231.3072 | 1.692 | 194.3402
|  | 1.3793 | 210.9023 | [8 empty] | 1.4322 | 203.4629 | 1.1566 | 231.2718 | 1.6925 | 194.0794
|  | 1.3794 | 210.8514 | [8 empty] | 1.4323 | 203.4068 | 1.1566 | 231.2365 | 1.6929 | 193.8179
|  | 1.3794 | 210.8005 | [8 empty] | 1.4324 | 203.3506 | 1.1567 | 231.2012 | 1.6933 | 193.5557
|  | 1.3795 | 210.7496 | [8 empty] | 1.4324 | 203.2944 | 1.1567 | 231.1659 | 1.6938 | 193.2929
|  | 1.3796 | 210.6986 | [8 empty] | 1.4325 | 203.2382 | 1.1568 | 231.1305 | 1.6942 | 193.0294
|  | 1.3796 | 210.6477 | [8 empty] | 1.4326 | 203.182 | 1.1568 | 231.0952 | 1.6946 | 192.7652
|  | 1.3797 | 210.5968 | [8 empty] | 1.4326 | 203.1258 | 1.1569 | 231.06 | 1.6951 | 192.5006
|  | 1.3797 | 210.5458 | [8 empty] | 1.4327 | 203.0696 | 1.1569 | 231.0247 | 1.6955 | 192.2352
|  | 1.3798 | 210.4949 | [8 empty] | 1.4328 | 203.0134 | 1.157 | 230.9894 | 1.6959 | 191.9692
|  | 1.3799 | 210.4439 | [8 empty] | 1.4328 | 202.9571 | 1.157 | 230.9541 | 1.6964 | 191.7024
|  | 1.3799 | 210.393 | [8 empty] | 1.4329 | 202.9009 | 1.1571 | 230.9189 | 1.6968 | 191.4349
|  | 1.38 | 210.342 | [8 empty] | 1.433 | 202.8447 | 1.1571 | 230.8836 | 1.6973 | 191.1667
|  | 1.3801 | 210.291 | [8 empty] | 1.433 | 202.7884 | 1.1572 | 230.8484 | 1.6977 | 190.898
|  | 1.3801 | 210.2401 | [8 empty] | 1.4331 | 202.7321 | 1.1572 | 230.8131 | 1.6981 | 190.6286
|  | 1.3802 | 210.1891 | [8 empty] | 1.4332 | 202.6759 | 1.1573 | 230.7779 | 1.6986 | 190.3584
|  | 1.3803 | 210.1381 | [8 empty] | 1.4332 | 202.6196 | 1.1573 | 230.7427 | 1.699 | 190.0874
|  | 1.3803 | 210.0871 | [8 empty] | 1.4333 | 202.5633 | 1.1574 | 230.7075 | 1.6994 | 189.8157
|  | 1.3804 | 210.0362 | [8 empty] | 1.4334 | 202.507 | 1.1574 | 230.6723 | 1.6999 | 189.5433
|  | 1.3804 | 209.9852 | [8 empty] | 1.4334 | 202.4507 | 1.1574 | 230.6371 | 1.7003 | 189.2703
|  | 1.3805 | 209.9342 | [8 empty] | 1.4335 | 202.3944 | 1.1575 | 230.6019 | 1.7007 | 188.9965
|  | 1.3806 | 209.8832 | [8 empty] | 1.4336 | 202.3381 | 1.1575 | 230.5667 | 1.7012 | 188.7218
|  | 1.3806 | 209.8322 | [8 empty] | 1.4336 | 202.2818 | 1.1576 | 230.5316 | 1.7016 | 188.4464
|  | 1.3807 | 209.7812 | [8 empty] | 1.4337 | 202.2255 | 1.1576 | 230.4964 | 1.7021 | 188.1702
|  | 1.3808 | 209.7302 | [8 empty] | 1.4338 | 202.1692 | 1.1577 | 230.4612 | 1.7025 | 187.8932
|  | 1.3808 | 209.6791 | [8 empty] | 1.4338 | 202.1128 | 1.1577 | 230.4261 | 1.7029 | 187.6156
|  | 1.3809 | 209.6281 | [8 empty] | 1.4339 | 202.0565 | 1.1578 | 230.391 | 1.7034 | 187.3371
|  | 1.3809 | 209.5771 | [8 empty] | 1.434 | 202.0001 | 1.1578 | 230.3558 | 1.7038 | 187.0578
|  | 1.381 | 209.5261 | [8 empty] | 1.434 | 201.9438 | 1.1579 | 230.3207 | 1.7042 | 186.7775
|  | 1.3811 | 209.475 | [8 empty] | 1.4341 | 201.8874 | 1.1579 | 230.2856 | 1.7047 | 186.4965
|  | 1.3811 | 209.424 | [8 empty] | 1.4342 | 201.8311 | 1.158 | 230.2505 | 1.7051 | 186.2146
|  | 1.3812 | 209.373 | [8 empty] | 1.4342 | 201.7747 | 1.158 | 230.2154 | 1.7056 | 185.9318
|  | 1.3813 | 209.3219 | [8 empty] | 1.4343 | 201.7183 | 1.1581 | 230.1804 | 1.706 | 185.6483
|  | 1.3813 | 209.2709 | [8 empty] | 1.4344 | 201.6619 | 1.1581 | 230.1453 | 1.7064 | 185.3639
|  | 1.3814 | 209.2198 | [8 empty] | 1.4344 | 201.6055 | 1.1582 | 230.1102 | 1.7069 | 185.0785
|  | 1.3814 | 209.1687 | [8 empty] | 1.4345 | 201.5491 | 1.1582 | 230.0752 | 1.7073 | 184.7923
|  | 1.3815 | 209.1177 | [8 empty] | 1.4346 | 201.4927 | 1.1583 | 230.0401 | 1.7077 | 184.5051
|  | 1.3816 | 209.0666 | [8 empty] | 1.4346 | 201.4363 | 1.1583 | 230.0051 | 1.7082 | 184.217
|  | 1.3816 | 209.0155 | [8 empty] | 1.4347 | 201.3799 | 1.1584 | 229.97 | 1.7086 | 183.9278
|  | 1.3817 | 208.9644 | [8 empty] | 1.4348 | 201.3235 | 1.1584 | 229.935 | 1.709 | 183.6379
|  | 1.3818 | 208.9134 | [8 empty] | 1.4348 | 201.2671 | 1.1585 | 229.9 | 1.7095 | 183.347
|  | 1.3818 | 208.8623 | [8 empty] | 1.4349 | 201.2106 | 1.1585 | 229.865 | 1.7099 | 183.0552
|  | 1.3819 | 208.8112 | [8 empty] | 1.435 | 201.1542 | 1.1585 | 229.83 | 1.7104 | 182.7624
|  | 1.382 | 208.7601 | [8 empty] | 1.435 | 201.0978 | 1.1586 | 229.795 | 1.7108 | 182.4685
|  | 1.382 | 208.709 | [8 empty] | 1.4351 | 201.0413 | 1.1586 | 229.76 | 1.7112 | 182.1734
|  | 1.3821 | 208.6579 | [8 empty] | 1.4352 | 200.9849 | 1.1587 | 229.725 | 1.7117 | 181.8773
|  | 1.3821 | 208.6067 | [8 empty] | 1.4352 | 200.9284 | 1.1587 | 229.6901 | 1.7121 | 181.5802
|  | 1.3822 | 208.5556 | [8 empty] | 1.4353 | 200.8719 | 1.1588 | 229.6551 | 1.7125 | 181.2824
|  | 1.3823 | 208.5045 | [8 empty] | 1.4354 | 200.8155 | 1.1588 | 229.6202 | 1.713 | 180.9834
|  | 1.3823 | 208.4534 | [8 empty] | 1.4354 | 200.759 | 1.1589 | 229.5852 | 1.7134 | 180.6832
|  | 1.3824 | 208.4022 | [8 empty] | 1.4355 | 200.7025 | 1.1589 | 229.5503 | 1.7138 | 180.3819
|  | 1.3825 | 208.3511 | [8 empty] | 1.4356 | 200.646 | 1.159 | 229.5154 | 1.7143 | 180.0793
|  | 1.3825 | 208.2999 | [8 empty] | 1.4356 | 200.5895 | 1.159 | 229.4805 | 1.7147 | 179.7755
|  | 1.3826 | 208.2488 | [8 empty] | 1.4357 | 200.533 | 1.1591 | 229.4456 | 1.7152 | 179.4707
|  | 1.3826 | 208.1976 | [8 empty] | 1.4358 | 200.4766 | 1.1591 | 229.4107 | 1.7156 | 179.1646
|  | 1.3827 | 208.1465 | [8 empty] | 1.4358 | 200.4201 | 1.1592 | 229.3758 | 1.716 | 178.8573
|  | 1.3828 | 208.0953 | [8 empty] | 1.4359 | 200.3635 | 1.1592 | 229.3409 | 1.7165 | 178.549
|  | 1.3828 | 208.0441 | [8 empty] | 1.436 | 200.307 | 1.1593 | 229.306 | 1.7169 | 178.2395
|  | 1.3829 | 207.9929 | [8 empty] | 1.436 | 200.2505 | 1.1593 | 229.2712 | 1.7173 | 177.9286
|  | 1.383 | 207.9417 | [8 empty] | 1.4361 | 200.194 | 1.1594 | 229.2363 | 1.7178 | 177.6166
|  | 1.383 | 207.8906 | [8 empty] | 1.4362 | 200.1375 | 1.1594 | 229.2015 | 1.7182 | 177.3032
|  | 1.3831 | 207.8394 | [8 empty] | 1.4362 | 200.081 | 1.1595 | 229.1666 | 1.7186 | 176.9884
|  | 1.3831 | 207.7882 | [8 empty] | 1.4363 | 200.0244 | 1.1595 | 229.1318 | 1.7191 | 176.6722
|  | 1.3832 | 207.737 | [8 empty] | 1.4364 | 199.9679 | 1.1596 | 229.097 | 1.7195 | 176.3545
|  | 1.3833 | 207.6857 | [8 empty] | 1.4364 | 199.9114 | 1.1596 | 229.0622 | 1.72 | 176.0357
|  | 1.3833 | 207.6345 | [8 empty] | 1.4365 | 199.8548 | 1.1596 | 229.0273 | 1.7204 | 175.7153
|  | 1.3834 | 207.5833 | [8 empty] | 1.4366 | 199.7983 | 1.1597 | 228.9925 | 1.7208 | 175.3939
|  | 1.3835 | 207.5321 | [8 empty] | 1.4366 | 199.7417 | 1.1597 | 228.9578 | 1.7213 | 175.0709
|  | 1.3835 | 207.4808 | [8 empty] | 1.4367 | 199.6852 | 1.1598 | 228.923 | 1.7217 | 174.7464
|  | 1.3836 | 207.4296 | [8 empty] | 1.4368 | 199.6286 | 1.1598 | 228.8882 | 1.7221 | 174.4204
|  | 1.3837 | 207.3784 | [8 empty] | 1.4368 | 199.5721 | 1.1599 | 228.8534 | 1.7226 | 174.093
|  | 1.3837 | 207.3271 | [8 empty] | 1.4369 | 199.5155 | 1.1599 | 228.8187 | 1.723 | 173.7639
|  | 1.3838 | 207.2758 | [8 empty] | 1.437 | 199.459 | 1.16 | 228.7839 | 1.7235 | 173.4332
|  | 1.3838 | 207.2246 | [8 empty] | 1.437 | 199.4024 | 1.16 | 228.7492 | 1.7239 | 173.1008
|  | 1.3839 | 207.1733 | [8 empty] | 1.4371 | 199.3458 | 1.1601 | 228.7145 | 1.7243 | 172.7667
|  | 1.384 | 207.122 | [8 empty] | 1.4372 | 199.2893 | 1.1601 | 228.6797 | 1.7248 | 172.4309
|  | 1.384 | 207.0708 | [8 empty] | 1.4372 | 199.2327 | 1.1602 | 228.645 | 1.7252 | 172.0936
|  | 1.3841 | 207.0195 | [8 empty] | 1.4373 | 199.1761 | 1.1602 | 228.6103 | 1.7256 | 171.7544
|  | 1.3842 | 206.9682 | [8 empty] | 1.4374 | 199.1196 | 1.1603 | 228.5756 | 1.7261 | 171.414
|  | 1.3842 | 206.9169 | [8 empty] | 1.4374 | 199.063 | 1.1603 | 228.5409 | 1.7265 | 171.0718
|  | 1.3843 | 206.8656 | [8 empty] | 1.4375 | 199.0064 | 1.1604 | 228.5063 | 1.7269 | 170.7277
|  | 1.3843 | 206.8143 | [8 empty] | 1.4376 | 198.9498 | 1.1604 | 228.4716 | 1.7274 | 170.3817
|  | 1.3844 | 206.763 | [8 empty] | 1.4376 | 198.8933 | 1.1605 | 228.4369 | 1.7278 | 170.034
|  | 1.3845 | 206.7116 | [8 empty] | 1.4377 | 198.8367 | 1.1605 | 228.4023 | 1.7283 | 169.6844
|  | 1.3845 | 206.6603 | [8 empty] | 1.4378 | 198.7801 | 1.1606 | 228.3676 | 1.7287 | 169.3328
|  | 1.3846 | 206.609 | [8 empty] | 1.4378 | 198.7235 | 1.1606 | 228.333 | 1.7291 | 168.9793
|  | 1.3847 | 206.5576 | [8 empty] | 1.4379 | 198.6669 | 1.1607 | 228.2983 | 1.7296 | 168.6236
|  | 1.3847 | 206.5063 | [8 empty] | 1.438 | 198.6104 | 1.1607 | 228.2637 | 1.73 | 168.2663
|  | 1.3848 | 206.4549 | [8 empty] | 1.438 | 198.5538 | 1.1607 | 228.2291 | 1.7304 | 167.9069
|  | 1.3848 | 206.4036 | [8 empty] | 1.4381 | 198.4972 | 1.1608 | 228.1945 | 1.7309 | 167.5453
|  | 1.3849 | 206.3522 | [8 empty] | 1.4382 | 198.4406 | 1.1608 | 228.1599 | 1.7313 | 167.1816
|  | 1.385 | 206.3008 | [8 empty] | 1.4382 | 198.384 | 1.1609 | 228.1253 | 1.7317 | 166.8158
|  | 1.385 | 206.2495 | [8 empty] | 1.4383 | 198.3274 | 1.1609 | 228.0907 | 1.7322 | 166.4485
|  | 1.3851 | 206.1981 | [8 empty] | 1.4384 | 198.2708 | 1.161 | 228.0562 | 1.7326 | 166.0789
|  | 1.3852 | 206.1467 | [8 empty] | 1.4384 | 198.2143 | 1.161 | 228.0216 | 1.7331 | 165.7076
|  | 1.3852 | 206.0953 | [8 empty] | 1.4385 | 198.1577 | 1.1611 | 227.987 | 1.7335 | 165.3341
|  | 1.3853 | 206.0439 | [8 empty] | 1.4386 | 198.1011 | 1.1611 | 227.9525 | 1.7339 | 164.9584
|  | 1.3854 | 205.9925 | [8 empty] | 1.4386 | 198.0445 | 1.1612 | 227.9179 | 1.7344 | 164.5805
|  | 1.3854 | 205.9411 | [8 empty] | 1.4387 | 197.9879 | 1.1612 | 227.8834 | 1.7348 | 164.2005
|  | 1.3855 | 205.8896 | [8 empty] | 1.4388 | 197.9313 | 1.1613 | 227.8489 | 1.7352 | 163.8183
|  | 1.3855 | 205.8382 | [8 empty] | 1.4388 | 197.8748 | 1.1613 | 227.8144 | 1.7357 | 163.4345
|  | 1.3856 | 205.7868 | [8 empty] | 1.4389 | 197.8182 | 1.1614 | 227.7799 | 1.7361 | 163.0493
|  | 1.3857 | 205.7353 | [8 empty] | 1.439 | 197.7616 | 1.1614 | 227.7454 | 1.7365 | 162.6622
|  | 1.3857 | 205.6839 | [8 empty] | 1.439 | 197.705 | 1.1615 | 227.7109 | 1.737 | 162.2731
|  | 1.3858 | 205.6324 | [8 empty] | 1.4391 | 197.6485 | 1.1615 | 227.6764 | 1.7374 | 161.8822
|  | 1.3859 | 205.581 | [8 empty] | 1.4392 | 197.5919 | 1.1616 | 227.6419 | 1.7379 | 161.4897
|  | 1.3859 | 205.5295 | [8 empty] | 1.4392 | 197.5353 | 1.1616 | 227.6074 | 1.7383 | 161.0964
|  | 1.386 | 205.478 | [8 empty] | 1.4393 | 197.4787 | 1.1617 | 227.573 | 1.7387 | 160.7018
|  | 1.386 | 205.4266 | [8 empty] | 1.4394 | 197.4222 | 1.1617 | 227.5385 | 1.7392 | 160.3065
|  | 1.3861 | 205.3751 | [8 empty] | 1.4394 | 197.3656 | 1.1618 | 227.5041 | 1.7396 | 159.9111
|  | 1.3862 | 205.3236 | [8 empty] | 1.4395 | 197.309 | 1.1618 | 227.4697 | 1.74 | 159.5159
|  | 1.3862 | 205.2721 | [8 empty] | 1.4396 | 197.2525 | 1.1618 | 227.4352 | 1.7405 | 159.125
|  | 1.3863 | 205.2206 | [8 empty] | 1.4396 | 197.1959 | 1.1619 | 227.4008 | 1.7409 | 158.7341
|  | 1.3864 | 205.1691 | [8 empty] | 1.4397 | 197.1394 | 1.1619 | 227.3664 | 1.7414 | 158.3413
|  | 1.3864 | 205.1176 | [8 empty] | 1.4398 | 197.0828 | 1.162 | 227.332 | 1.7418 | 157.9505
|  | 1.3865 | 205.066 | [8 empty] | 1.4398 | 197.0263 | 1.162 | 227.2976 | 1.7422 | 157.5627
|  | 1.3865 | 205.0145 | [8 empty] | 1.4399 | 196.9697 | 1.1621 | 227.2632 | 1.7427 | 157.1786
|  | 1.3866 | 204.963 | [8 empty] | 1.44 | 196.9132 | 1.1621 | 227.2288 | 1.7431 | 156.799
|  | 1.3867 | 204.9114 | [8 empty] | 1.44 | 196.8567 | 1.1622 | 227.1945 | 1.7435 | 156.425
|  | 1.3867 | 204.8599 | [8 empty] | 1.4401 | 196.8001 | 1.1622 | 227.1601 | 1.744 | 156.0584
|  | 1.3868 | 204.8083 | [8 empty] | 1.4402 | 196.7436 | 1.1623 | 227.1258 | 1.7444 | 155.6999
|  | 1.3869 | 204.7567 | [8 empty] | 1.4402 | 196.6871 | 1.1623 | 227.0914 | 1.7448 | 155.3503
|  | 1.3869 | 204.7052 | [8 empty] | 1.4403 | 196.6306 | 1.1624 | 227.0571 | 1.7453 | 155.0117
|  | 1.387 | 204.6536 | [8 empty] | 1.4404 | 196.5741 | 1.1624 | 227.0227 | 1.7457 | 154.6854
|  | 1.3871 | 204.602 | [8 empty] | 1.4404 | 196.5176 | 1.1625 | 226.9884 | 1.7462 | 154.3725
|  | 1.3871 | 204.5504 | [8 empty] | 1.4405 | 196.4611 | 1.1625 | 226.9541 | 1.7466 | 154.0751
|  | 1.3872 | 204.4988 | [8 empty] | 1.4406 | 196.4046 | 1.1626 | 226.9198 | 1.747 | 153.7938
|  | 1.3872 | 204.4472 | [8 empty] | 1.4406 | 196.3481 | 1.1626 | 226.8855 | 1.7475 | 153.5269
|  | 1.3873 | 204.3956 | [8 empty] | 1.4407 | 196.2916 | 1.1627 | 226.8512 | 1.7479 | 153.2797
|  | 1.3874 | 204.3439 | [8 empty] | 1.4408 | 196.2351 | 1.1627 | 226.8169 | 1.7483 | 153.0547
|  | 1.3874 | 204.2923 | [8 empty] | 1.4408 | 196.1786 | 1.1628 | 226.7827 | 1.7488 | 152.8489
|  | 1.3875 | 204.2407 | [8 empty] | 1.4409 | 196.1222 | 1.1628 | 226.7484 | 1.7492 | 152.612
|  | 1.3876 | 204.189 | [8 empty] | 1.441 | 196.0657 | 1.1629 | 226.7142 | 1.7496 | 152.3916
|  | 1.3876 | 204.1374 | [8 empty] | 1.441 | 196.0093 | 1.1629 | 226.6799 | 1.7501 | 152.1635
|  | 1.3877 | 204.0857 | [8 empty] | 1.4411 | 195.9528 | 1.1629 | 226.6457 | 1.7505 | 151.94
|  | 1.3877 | 204.0341 | [8 empty] | 1.4411 | 195.8964 | 1.163 | 226.6114 | 1.751 | 151.7149
|  | 1.3878 | 203.9824 | [8 empty] | 1.4412 | 195.84 | 1.163 | 226.5772 | 1.7514 | 151.4913
|  | 1.3879 | 203.9307 | [8 empty] | 1.4413 | 195.7835 | 1.1631 | 226.543 | 1.7518 | 151.2677
|  | 1.3879 | 203.879 | [8 empty] | 1.4413 | 195.7271 | 1.1631 | 226.5088 | 1.7523 | 151.0447
|  | 1.388 | 203.8273 | [8 empty] | 1.4414 | 195.6707 | 1.1632 | 226.4746 | 1.7527 | 150.8221
|  | 1.3881 | 203.7756 | [8 empty] | 1.4415 | 195.6143 | 1.1632 | 226.4404 | 1.7531 | 150.6
|  | 1.3881 | 203.7239 | [8 empty] | 1.4415 | 195.5579 | 1.1633 | 226.4062 | 1.7536 | 150.3784
|  | 1.3882 | 203.6722 | [8 empty] | 1.4416 | 195.5015 | 1.1633 | 226.372 | 1.754 | 150.1571
|  | 1.3882 | 203.6205 | [8 empty] | 1.4417 | 195.4452 | 1.1634 | 226.3379 | 1.7544 | 149.9364
|  | 1.3883 | 203.5687 | [8 empty] | 1.4417 | 195.3888 | 1.1634 | 226.3037 | 1.7549 | 149.7161
|  | 1.3884 | 203.517 | [8 empty] | 1.4418 | 195.3325 | 1.1635 | 226.2696 | 1.7553 | 149.4962
|  | 1.3884 | 203.4653 | [8 empty] | 1.4419 | 195.2761 | 1.1635 | 226.2354 | 1.7558 | 149.2768
|  | 1.3885 | 203.4135 | [8 empty] | 1.4419 | 195.2198 | 1.1636 | 226.2013 | 1.7562 | 149.0578
|  | 1.3886 | 203.3618 | [8 empty] | 1.442 | 195.1634 | 1.1636 | 226.1672 | 1.7566 | 148.8393
|  | 1.3886 | 203.31 | [8 empty] | 1.4421 | 195.1071 | 1.1637 | 226.133 | 1.7571 | 148.6212
|  | 1.3887 | 203.2582 | [8 empty] | 1.4421 | 195.0508 | 1.1637 | 226.0989 | 1.7575 | 148.4036
|  | 1.3888 | 203.2064 | [8 empty] | 1.4422 | 194.9945 | 1.1638 | 226.0648 | 1.7579 | 148.1863
|  | 1.3888 | 203.1546 | [8 empty] | 1.4423 | 194.9383 | 1.1638 | 226.0307 | 1.7584 | 147.9696
|  | 1.3889 | 203.1028 | [8 empty] | 1.4423 | 194.882 | 1.1639 | 225.9967 | 1.7588 | 147.7532
|  | 1.3889 | 203.051 | [8 empty] | 1.4424 | 194.8257 | 1.1639 | 225.9626 | 1.7593 | 147.5373
|  | 1.389 | 202.9992 | [8 empty] | 1.4425 | 194.7695 | 1.164 | 225.9285 | 1.7597 | 147.3219
|  | 1.3891 | 202.9474 | [8 empty] | 1.4425 | 194.7132 | 1.164 | 225.8944 | 1.7601 | 147.1068
|  | 1.3891 | 202.8956 | [8 empty] | 1.4426 | 194.657 | 1.164 | 225.8604 | 1.7606 | 146.8922
|  | 1.3892 | 202.8437 | [8 empty] | 1.4427 | 194.6008 | 1.1641 | 225.8263 | 1.761 | 146.6781
|  | 1.3893 | 202.7919 | [8 empty] | 1.4427 | 194.5446 | 1.1641 | 225.7923 | 1.7614 | 146.4643
|  | 1.3893 | 202.74 | [8 empty] | 1.4428 | 194.4884 | 1.1642 | 225.7583 | 1.7619 | 146.251
|  | 1.3894 | 202.6882 | [8 empty] | 1.4429 | 194.4322 | 1.1642 | 225.7243 | 1.7623 | 146.0381
|  | 1.3894 | 202.6363 | [8 empty] | 1.4429 | 194.3761 | 1.1643 | 225.6902 | 1.7627 | 145.8256
|  | 1.3895 | 202.5845 | [8 empty] | 1.443 | 194.3199 | 1.1643 | 225.6562 | 1.7632 | 145.6135
|  | 1.3896 | 202.5326 | [8 empty] | 1.4431 | 194.2638 | 1.1644 | 225.6222 | 1.7636 | 145.4019
|  | 1.3896 | 202.4807 | [8 empty] | 1.4431 | 194.2077 | 1.1644 | 225.5883 | 1.7641 | 145.1907
|  | 1.3897 | 202.4288 | [8 empty] | 1.4432 | 194.1516 | 1.1645 | 225.5543 | 1.7645 | 144.9799
|  | 1.3898 | 202.3769 | [8 empty] | 1.4433 | 194.0955 | 1.1645 | 225.5203 | 1.7649 | 144.7695
|  | 1.3898 | 202.325 | [8 empty] | 1.4433 | 194.0394 | 1.1646 | 225.4863 | 1.7654 | 144.5596
|  | 1.3899 | 202.2731 | [8 empty] | 1.4434 | 193.9833 | 1.1646 | 225.4524 | 1.7658 | 144.35
|  | 1.3899 | 202.2211 | [8 empty] | 1.4435 | 193.9273 | 1.1647 | 225.4184 | 1.7662 | 144.1409
|  | 1.39 | 202.1692 | [8 empty] | 1.4435 | 193.8713 | 1.1647 | 225.3845 | 1.7667 | 143.9321
|  | 1.3901 | 202.1172 | [8 empty] | 1.4436 | 193.8152 | 1.1648 | 225.3505 | 1.7671 | 143.7238
|  | 1.3901 | 202.0653 | [8 empty] | 1.4437 | 193.7592 | 1.1648 | 225.3166 | 1.7675 | 143.5159
|  | 1.3902 | 202.0133 | [8 empty] | 1.4437 | 193.7033 | 1.1649 | 225.2827 | 1.768 | 143.3084
|  | 1.3903 | 201.9614 | [8 empty] | 1.4438 | 193.6473 | 1.1649 | 225.2488 | 1.7684 | 143.1013
|  | 1.3903 | 201.9094 | [8 empty] | 1.4439 | 193.5913 | 1.165 | 225.2149 | 1.7689 | 142.8946
|  | 1.3904 | 201.8574 | [8 empty] | 1.4439 | 193.5354 | 1.165 | 225.181 | 1.7693 | 142.6883
|  | 1.3905 | 201.8054 | [8 empty] | 1.444 | 193.4795 | 1.1651 | 225.1471 | 1.7697 | 142.4824
|  | 1.3905 | 201.7534 | [8 empty] | 1.4441 | 193.4236 | 1.1651 | 225.1132 | 1.7702 | 142.2769
|  | 1.3906 | 201.7014 | [8 empty] | 1.4441 | 193.3677 | 1.1651 | 225.0794 | 1.7706 | 142.0719
|  | 1.3906 | 201.6494 | [8 empty] | 1.4442 | 193.3118 | 1.1652 | 225.0455 | 1.771 | 141.8672
|  | 1.3907 | 201.5974 | [8 empty] | 1.4443 | 193.256 | 1.1652 | 225.0116 | 1.7715 | 141.6629
|  | 1.3908 | 201.5454 | [8 empty] | 1.4443 | 193.2002 | 1.1653 | 224.9778 | 1.7719 | 141.459
|  | 1.3908 | 201.4933 | [8 empty] | 1.4444 | 193.1443 | 1.1653 | 224.944 | 1.7723 | 141.2555
|  | 1.3909 | 201.4413 | [8 empty] | 1.4445 | 193.0886 | 1.1654 | 224.9101 | 1.7728 | 141.0524
|  | 1.391 | 201.3892 | [8 empty] | 1.4445 | 193.0328 | 1.1654 | 224.8763 | 1.7732 | 140.8496
|  | 1.391 | 201.3372 | [8 empty] | 1.4446 | 192.977 | 1.1655 | 224.8425 | 1.7737 | 140.6473
|  | 1.3911 | 201.2851 | [8 empty] | 1.4447 | 192.9213 | 1.1655 | 224.8087 | 1.7741 | 140.4454
|  | 1.3911 | 201.233 | [8 empty] | 1.4447 | 192.8656 | 1.1656 | 224.7749 | 1.7745 | 140.2438
|  | 1.3912 | 201.1809 | [8 empty] | 1.4448 | 192.8099 | 1.1656 | 224.7411 | 1.775 | 140.0426
|  | 1.3913 | 201.1288 | [8 empty] | 1.4449 | 192.7542 | 1.1657 | 224.7073 | 1.7754 | 139.8419
|  | 1.3913 | 201.0767 | [8 empty] | 1.4449 | 192.6986 | 1.1657 | 224.6735 | 1.7758 | 139.6415
|  | 1.3914 | 201.0246 | [8 empty] | 1.445 | 192.6429 | 1.1658 | 224.6398 | 1.7763 | 139.4414
|  | 1.3915 | 200.9725 | [8 empty] | 1.4451 | 192.5873 | 1.1658 | 224.606 | 1.7767 | 139.2418
|  | 1.3915 | 200.9204 | [8 empty] | 1.4451 | 192.5317 | 1.1659 | 224.5723 | 1.7772 | 139.0426
|  | 1.3916 | 200.8683 | [8 empty] | 1.4452 | 192.4762 | 1.1659 | 224.5385 | 1.7776 | 138.8437
|  | 1.3916 | 200.8161 | [8 empty] | 1.4453 | 192.4206 | 1.166 | 224.5048 | 1.778 | 138.6452
|  | 1.3917 | 200.764 | [8 empty] | 1.4453 | 192.3651 | 1.166 | 224.471 | 1.7785 | 138.4471
|  | 1.3918 | 200.7118 | [8 empty] | 1.4454 | 192.3096 | 1.1661 | 224.4373 | 1.7789 | 138.2493
|  | 1.3918 | 200.6597 | [8 empty] | 1.4455 | 192.2541 | 1.1661 | 224.4036 | 1.7793 | 138.0519
|  | 1.3919 | 200.6075 | [8 empty] | 1.4455 | 192.1987 | 1.1662 | 224.3699 | 1.7798 | 137.8549
|  | 1.392 | 200.5553 | [8 empty] | 1.4456 | 192.1433 | 1.1662 | 224.3362 | 1.7802 | 137.6583
|  | 1.392 | 200.5031 | [8 empty] | 1.4457 | 192.0878 | 1.1662 | 224.3025 | 1.7806 | 137.4621
|  | 1.3921 | 200.4509 | [8 empty] | 1.4457 | 192.0325 | 1.1695 | 223.8951 | 1.7811 | 137.2662
|  | 1.3922 | 200.3987 | [8 empty] | 1.4458 | 191.9771 | 1.1727 | 223.4867 | 1.7815 | 137.0707
|  | 1.3922 | 200.3465 | [8 empty] | 1.4459 | 191.9218 | 1.176 | 223.0772 | 1.782 | 136.8755
|  | 1.3923 | 200.2943 | [8 empty] | 1.4459 | 191.8665 | 1.1792 | 222.6667 | 1.7824 | 136.6807
|  | 1.3923 | 200.242 | [8 empty] | 1.446 | 191.8112 | 1.1825 | 222.2552 | 1.7828 | 136.4863
|  | 1.3924 | 200.1898 | [8 empty] | 1.4461 | 191.7559 | 1.1857 | 221.8426 | 1.7833 | 136.2923
|  | 1.3925 | 200.1376 | [8 empty] | 1.4461 | 191.7007 | 1.1889 | 221.4291 | 1.7837 | 136.0986
|  | 1.3925 | 200.0853 | [8 empty] | 1.4462 | 191.6455 | 1.1922 | 221.0144 | 1.7841 | 135.9052
|  | 1.3926 | 200.033 | [8 empty] | 1.4463 | 191.5903 | 1.1954 | 220.5988 | 1.7846 | 135.7123
|  | 1.3927 | 199.9808 | [8 empty] | 1.4463 | 191.5352 | 1.1987 | 220.1822 | 1.785 | 135.5197
|  | 1.3927 | 199.9285 | [8 empty] | 1.4464 | 191.48 | 1.2019 | 219.7645 | 1.7854 | 135.3274
|  | 1.3928 | 199.8762 | [8 empty] | 1.4465 | 191.4249 | 1.2051 | 219.3458 | 1.7859 | 135.1355
|  | 1.3928 | 199.8239 | [8 empty] | 1.4465 | 191.3699 | 1.2084 | 218.9261 | 1.7863 | 134.944
|  | 1.3929 | 199.7716 | [8 empty] | 1.4466 | 191.3148 | 1.2116 | 218.5054 | 1.7868 | 134.7528
|  | 1.393 | 199.7193 | [8 empty] | 1.4467 | 191.2598 | 1.2149 | 218.0837 | 1.7872 | 134.5619
|  | 1.393 | 199.667 | [8 empty] | 1.4467 | 191.2048 | 1.2181 | 217.6609 | 1.7876 | 134.3715
|  | 1.3931 | 199.6147 | [8 empty] | 1.4468 | 191.1499 | 1.2213 | 217.2372 | 1.7881 | 134.1813
|  | 1.3932 | 199.5623 | [8 empty] | 1.4469 | 191.0949 | 1.2246 | 216.8125 | 1.7885 | 133.9916
|  | 1.3932 | 199.51 | [8 empty] | 1.4469 | 191.04 | 1.2278 | 216.3868 | 1.7889 | 133.8021
|  | 1.3933 | 199.4576 | [8 empty] | 1.447 | 190.9852 | 1.2311 | 215.9601 | 1.7894 | 133.613
|  | 1.3933 | 199.4053 | [8 empty] | 1.4471 | 190.9303 | 1.2343 | 215.5324 | 1.7898 | 133.4243
|  | 1.3934 | 199.3529 | [8 empty] | 1.4471 | 190.8755 | 1.2375 | 215.1038 | 1.7902 | 133.2359
|  | 1.3935 | 199.3005 | [8 empty] | 1.4472 | 190.8207 | 1.2408 | 214.6741 | 1.7907 | 133.0479
|  | 1.3935 | 199.2481 | [8 empty] | 1.4473 | 190.766 | 1.244 | 214.2435 | 1.7911 | 132.8602
|  | 1.3936 | 199.1957 | [8 empty] | 1.4473 | 190.7113 | 1.2473 | 213.8119 | 1.7916 | 132.6728
|  | 1.3937 | 199.1433 | [8 empty] | 1.4474 | 190.6566 | 1.2505 | 213.3793 | 1.792 | 132.4858
|  | 1.3937 | 199.0909 | [8 empty] | 1.4475 | 190.6019 | 1.2537 | 212.9458 | 1.7924 | 132.2991
|  | 1.3938 | 199.0385 | [8 empty] | 1.4475 | 190.5473 | 1.257 | 212.5113 | 1.7929 | 132.1128
|  | 1.3939 | 198.9861 | [8 empty] | 1.4476 | 190.4927 | 1.2602 | 212.0759 | 1.7933 | 131.9268
|  | 1.3939 | 198.9337 | [8 empty] | 1.4477 | 190.4381 | 1.2635 | 211.6395 | 1.7937 | 131.7411
|  | 1.394 | 198.8812 | [8 empty] | 1.4477 | 190.3836 | 1.2667 | 211.2022 | 1.7942 | 131.5558
|  | 1.394 | 198.8288 | [8 empty] | 1.4478 | 190.3291 | 1.2699 | 210.7639 | 1.7946 | 131.3708
|  | 1.3941 | 198.7763 | [8 empty] | 1.4479 | 190.2746 | 1.2732 | 210.3247 | 1.7951 | 131.1862
|  | 1.3942 | 198.7239 | [8 empty] | 1.4479 | 190.2202 | 1.2764 | 209.8845 | 1.7955 | 131.0018
|  | 1.3942 | 198.6714 | [8 empty] | 1.448 | 190.1658 | 1.2797 | 209.4434 | 1.7959 | 130.8179
|  | 1.3943 | 198.6189 | [8 empty] | 1.4481 | 190.1114 | 1.2829 | 209.0014 | 1.7964 | 130.6342
|  | 1.3944 | 198.5664 | [8 empty] | 1.4481 | 190.0571 | 1.2861 | 208.5584 | 1.7968 | 130.4509
|  | 1.3944 | 198.5139 | [8 empty] | 1.4482 | 190.0028 | 1.2894 | 208.1146 | 1.7972 | 130.2679
|  | 1.3945 | 198.4614 | [8 empty] | 1.4483 | 189.9486 | 1.2926 | 207.6698 | 1.7977 | 130.0852
|  | 1.3945 | 198.4089 | [8 empty] | 1.4483 | 189.8943 | 1.2959 | 207.2241 | 1.7981 | 129.9028
|  | 1.3946 | 198.3564 | [8 empty] | 1.4484 | 189.8401 | 1.2991 | 206.7776 | 1.7985 | 129.7208
|  | 1.3947 | 198.3039 | [8 empty] | 1.4485 | 189.786 | 1.3023 | 206.3301 | 1.799 | 129.5391
|  | 1.3947 | 198.2513 | [8 empty] | 1.4485 | 189.7319 | 1.3056 | 205.8817 | 1.7994 | 129.3578
|  | 1.3948 | 198.1988 | [8 empty] | 1.4486 | 189.6778 | 1.3088 | 205.4324 | 1.7999 | 129.1767
|  | 1.3949 | 198.1462 | [8 empty] | 1.4487 | 189.6237 | 1.3121 | 204.9823 | 1.8003 | 128.996
|  | 1.3949 | 198.0937 | [8 empty] | 1.4487 | 189.5697 | 1.3153 | 204.5312 | 1.8007 | 128.8156
|  | 1.395 | 198.0411 | [8 empty] | 1.4488 | 189.5158 | 1.3185 | 204.0793 | 1.8012 | 128.6355
|  | 1.395 | 197.9885 | [8 empty] | 1.4489 | 189.4618 | 1.3218 | 203.6265 | 1.8016 | 128.4557
|  | 1.3951 | 197.9359 | [8 empty] | 1.4489 | 189.4079 | 1.325 | 203.1729 | 1.802 | 128.2763
|  | 1.3952 | 197.8833 | [8 empty] | 1.449 | 189.3541 | 1.3283 | 202.7184 | 1.8025 | 128.0972
|  | 1.3952 | 197.8307 | [8 empty] | 1.4491 | 189.3002 | 1.3315 | 202.263 | 1.8029 | 127.9184
|  | 1.3953 | 197.7781 | [8 empty] | 1.4491 | 189.2465 | 1.3347 | 201.8068 | 1.8033 | 127.7399
|  | 1.3954 | 197.7255 | [8 empty] | 1.4492 | 189.1927 | 1.338 | 201.3497 | 1.8038 | 127.5617
|  | 1.3954 | 197.6729 | [8 empty] | 1.4493 | 189.139 | 1.3412 | 200.8918 | 1.8042 | 127.3838
|  | 1.3955 | 197.6203 | [8 empty] | 1.4493 | 189.0853 | 1.3445 | 200.433 | 1.8047 | 127.2062
|  | 1.3956 | 197.5676 | [8 empty] | 1.4494 | 189.0317 | 1.3477 | 199.9734 | 1.8051 | 127.029
|  | 1.3956 | 197.515 | [8 empty] | 1.4495 | 188.9781 | 1.3509 | 199.513 | 1.8055 | 126.852
|  | 1.3957 | 197.4623 | [8 empty] | 1.4495 | 188.9246 | 1.3542 | 199.0518 | 1.806 | 126.6754
|  | 1.3957 | 197.4097 | [8 empty] | 1.4496 | 188.8711 | 1.3574 | 198.5897 | 1.8064 | 126.4991
|  | 1.3958 | 197.357 | [8 empty] | 1.4497 | 188.8176 | 1.3607 | 198.1269 | 1.8068 | 126.3231
|  | 1.3959 | 197.3043 | [8 empty] | 1.4497 | 188.7642 | 1.3639 | 197.6632 | 1.8073 | 126.1474
|  | 1.3959 | 197.2516 | [8 empty] | 1.4498 | 188.7108 | 1.3672 | 197.1987 | 1.8077 | 125.972
|  | 1.396 | 197.1989 | [8 empty] | 1.4499 | 188.6574 | 1.3704 | 196.7334 | 1.8081 | 125.7969
|  | 1.3961 | 197.1462 | [8 empty] | 1.4499 | 188.6041 | 1.3736 | 196.2674 | 1.8086 | 125.6221
|  | 1.3961 | 197.0935 | [8 empty] | 1.45 | 188.5509 | 1.3769 | 195.8005 | 1.809 | 125.4476
|  | 1.3962 | 197.0408 | [8 empty] | 1.4501 | 188.4976 | 1.3801 | 195.3329 | 1.8095 | 125.2734
|  | 1.3962 | 196.9881 | [8 empty] | 1.4501 | 188.4445 | 1.3834 | 194.8644 | 1.8099 | 125.0995
|  | 1.3963 | 196.9354 | [8 empty] | 1.4502 | 188.3913 | 1.3866 | 194.3952 | 1.8103 | 124.9259
|  | 1.3964 | 196.8826 | [8 empty] | 1.4503 | 188.3382 | 1.3898 | 193.9253 | 1.8108 | 124.7527
|  | 1.3964 | 196.8299 | [8 empty] | 1.4503 | 188.2852 | 1.3931 | 193.4546 | 1.8112 | 124.5797
|  | 1.3965 | 196.7771 | [8 empty] | 1.4504 | 188.2322 | 1.3963 | 192.9831 | 1.8116 | 124.407
|  | 1.3966 | 196.7244 | [8 empty] | 1.4505 | 188.1792 | 1.3996 | 192.5108 | 1.8121 | 124.2346
|  | 1.3966 | 196.6716 | [8 empty] | 1.4505 | 188.1263 | 1.4028 | 192.0379 | 1.8125 | 124.0625
|  | 1.3967 | 196.6188 | [8 empty] | 1.4506 | 188.0734 | 1.406 | 191.5641 | 1.813 | 123.8907
|  | 1.3967 | 196.566 | [8 empty] | 1.4507 | 188.0206 | 1.4093 | 191.0897 | 1.8134 | 123.7192
|  | 1.3968 | 196.5132 | [8 empty] | 1.4507 | 187.9678 | 1.4125 | 190.6145 | 1.8138 | 123.5479
|  | 1.3969 | 196.4604 | [8 empty] | 1.4508 | 187.9151 | 1.4158 | 190.1385 | 1.8143 | 123.377
|  | 1.3969 | 196.4076 | [8 empty] | 1.4509 | 187.8624 | 1.419 | 189.6619 | 1.8147 | 123.2064
|  | 1.397 | 196.3548 | [8 empty] | 1.4509 | 187.8097 | 1.4222 | 189.1845 | 1.8151 | 123.036
|  | 1.3971 | 196.302 | [8 empty] | 1.451 | 187.7571 | 1.4255 | 188.7065 | 1.8156 | 122.866
|  | 1.3971 | 196.2492 | [8 empty] | 1.4511 | 187.7046 | 1.4287 | 188.2277 | 1.816 | 122.6962
|  | 1.3972 | 196.1963 | [8 empty] | 1.4511 | 187.652 | 1.432 | 187.7482 | 1.8164 | 122.5267
|  | 1.3973 | 196.1435 | [8 empty] | 1.4512 | 187.5996 | 1.4352 | 187.268 | 1.8169 | 122.3575
|  | 1.3973 | 196.0907 | [8 empty] | 1.4513 | 187.5471 | 1.4384 | 186.7872 | 1.8173 | 122.1886
|  | 1.3974 | 196.0378 | [8 empty] | 1.4513 | 187.4948 | 1.4417 | 186.3056 | 1.8178 | 122.02
|  | 1.3974 | 195.9849 | [8 empty] | 1.4514 | 187.4424 | 1.4449 | 185.8234 | 1.8182 | 121.8517
|  | 1.3975 | 195.9321 | [8 empty] | 1.4515 | 187.3902 | 1.4482 | 185.3405 | 1.8186 | 121.6836
|  | 1.3976 | 195.8792 | [8 empty] | 1.4515 | 187.3379 | 1.4514 | 184.8569 | 1.8191 | 121.5159
|  | 1.3976 | 195.8263 | [8 empty] | 1.4516 | 187.2857 | 1.4546 | 184.3727 | 1.8195 | 121.3484
|  | 1.3977 | 195.7734 | [8 empty] | 1.4517 | 187.2336 | 1.4579 | 183.8878 | 1.8199 | 121.1812
|  | 1.3978 | 195.7205 | [8 empty] | 1.4517 | 187.1815 | 1.4611 | 183.4022 | 1.8204 | 121.0142
|  | 1.3978 | 195.6676 | [8 empty] | 1.4518 | 187.1295 | 1.4644 | 182.916 | 1.8208 | 120.8476
|  | 1.3979 | 195.6147 | [8 empty] | 1.4519 | 187.0775 | 1.4676 | 182.4292 | 1.8212 | 120.6812
|  | 1.3979 | 195.5618 | [8 empty] | 1.4519 | 187.0255 | 1.4708 | 181.9417 | 1.8217 | 120.5151
|  | 1.398 | 195.5089 | [8 empty] | 1.452 | 186.9736 | 1.4741 | 181.4536 | 1.8221 | 120.3493
|  | 1.3981 | 195.456 | [8 empty] | 1.4521 | 186.9218 | 1.4773 | 180.9648 | 1.8226 | 120.1838
|  | 1.3981 | 195.403 | [8 empty] | 1.4521 | 186.87 | 1.4806 | 180.4754 | 1.823 | 120.0185
|  | 1.3982 | 195.3501 | [8 empty] | 1.4522 | 186.8182 | 1.4838 | 179.9854 | 1.8234 | 119.8535
|  | 1.3983 | 195.2971 | [8 empty] | 1.4523 | 186.7665 | 1.487 | 179.4948 | 1.8239 | 119.6888
|  | 1.3983 | 195.2442 | [8 empty] | 1.4523 | 186.7149 | 1.4903 | 179.0036 | 1.8243 | 119.5243
|  | 1.4016 | 194.5806 | [8 empty] | 1.4524 | 186.6633 | 1.4935 | 178.5118 | 1.8247 | 119.3602
|  | 1.4048 | 193.9167 | [8 empty] | 1.4525 | 186.6118 | 1.4968 | 178.0194 | 1.8252 | 119.1962
|  | 1.408 | 193.2525 | [8 empty] | 1.4525 | 186.5603 | 1.5 | 177.5264 | 1.8256 | 119.0326
|  | 1.4113 | 192.5879 | [8 empty] | 1.4526 | 186.5088 | 1.5032 | 177.0328 | 1.826 | 118.8692
|  | 1.4145 | 191.923 | [8 empty] | 1.4527 | 186.4574 | 1.5065 | 176.5386 | 1.8265 | 118.7062
|  | 1.4178 | 191.2578 | [8 empty] | 1.4527 | 186.4061 | 1.5097 | 176.0438 | 1.8269 | 118.5433
|  | 1.421 | 190.5923 | [8 empty] | 1.4528 | 186.3548 | 1.513 | 175.5485 | 1.8274 | 118.3808
|  | 1.4242 | 189.9264 | [8 empty] | 1.4529 | 186.3035 | 1.5162 | 175.0526 | 1.8278 | 118.2185
|  | 1.4275 | 189.2602 | [8 empty] | 1.4529 | 186.2523 | 1.5194 | 174.5562 | 1.8282 | 118.0564
|  | 1.4307 | 188.5937 | [8 empty] | 1.453 | 186.2012 | 1.5227 | 174.0592 | 1.8287 | 117.8947
|  | 1.434 | 187.9269 | [8 empty] | 1.4531 | 186.1501 | 1.5259 | 173.5616 | 1.8291 | 117.7332
|  | 1.4372 | 187.2598 | [8 empty] | 1.4531 | 186.0991 | 1.5292 | 173.0635 | 1.8295 | 117.5719
|  | 1.4404 | 186.5924 | [8 empty] | 1.4532 | 186.0481 | 1.5324 | 172.5648 | 1.83 | 117.411
|  | 1.4437 | 185.9247 | [8 empty] | 1.4533 | 185.9972 | 1.5357 | 172.0656 | 1.8304 | 117.2502
|  | 1.4469 | 185.2567 | [8 empty] | 1.4533 | 185.9463 | 1.5389 | 171.5659 | 1.8309 | 117.0898
|  | 1.4502 | 184.5884 | [8 empty] | 1.4534 | 185.8955 | 1.5421 | 171.0657 | 1.8313 | 116.9296
|  | 1.4534 | 183.9199 | [8 empty] | 1.4535 | 185.8447 | 1.5454 | 170.5649 | 1.8317 | 116.7697
|  | 1.4567 | 183.251 | [8 empty] | 1.4535 | 185.794 | 1.5486 | 170.0636 | 1.8322 | 116.61
|  | 1.4599 | 182.5819 | [8 empty] | 1.4536 | 185.7433 | 1.5519 | 169.5618 | 1.8326 | 116.4506
|  | 1.4631 | 181.9124 | [8 empty] | 1.4537 | 185.6927 | 1.5551 | 169.0595 | 1.833 | 116.2914
|  | 1.4664 | 181.2428 | [8 empty] | 1.4537 | 185.6421 | 1.5583 | 168.5567 | 1.8335 | 116.1325
|  | 1.4696 | 180.5728 | [8 empty] | 1.4538 | 185.5916 | 1.5616 | 168.0534 | 1.8339 | 115.9739
|  | 1.4729 | 179.9026 | [8 empty] | 1.4539 | 185.5412 | 1.5648 | 167.5496 | 1.8343 | 115.8155
|  | 1.4761 | 179.2321 | [8 empty] | 1.4539 | 185.4908 | 1.5681 | 167.0453 | 1.8348 | 115.6573
|  | 1.4793 | 178.5613 | [8 empty] | 1.454 | 185.4405 | 1.5713 | 166.5406 | 1.8352 | 115.4994
|  | 1.4826 | 177.8904 | [8 empty] | 1.4541 | 185.3902 | 1.5745 | 166.0353 | 1.8357 | 115.3418
|  | 1.4858 | 177.2191 | [8 empty] | 1.4541 | 185.3399 | 1.5778 | 165.5296 | 1.8361 | 115.1844
|  | 1.4891 | 176.5476 | [8 empty] | 1.4542 | 185.2898 | 1.581 | 165.0234 | 1.8365 | 115.0273
|  | 1.4923 | 175.8759 | [8 empty] | 1.4542 | 185.2396 | 1.5843 | 164.5168 | 1.837 | 114.8704
|  | 1.4955 | 175.2039 | [8 empty] | 1.4543 | 185.1896 | 1.5875 | 164.0097 | 1.8374 | 114.7138
|  | 1.4988 | 174.5317 | [8 empty] | 1.4544 | 185.1396 | 1.5907 | 163.5022 | 1.8378 | 114.5574
|  | 1.502 | 173.8593 | [8 empty] | 1.4544 | 185.0896 | 1.594 | 162.9942 | 1.8383 | 114.4012
|  | 1.5053 | 173.1867 | [8 empty] | 1.4545 | 185.0397 | 1.5972 | 162.4857 | 1.8387 | 114.2454
|  | 1.5085 | 172.5138 | [8 empty] | 1.4546 | 184.9899 | 1.6005 | 161.9769 | 1.8391 | 114.0897
|  | 1.5117 | 171.8407 | [8 empty] | 1.4546 | 184.9401 | 1.6037 | 161.4675 | 1.8396 | 113.9343
|  | 1.515 | 171.1674 | [8 empty] | 1.4547 | 184.8903 | 1.6069 | 160.9578 | 1.84 | 113.7792
|  | 1.5182 | 170.4939 | [8 empty] | 1.4548 | 184.8406 | 1.6102 | 160.4477 | 1.8405 | 113.6243
|  | 1.5215 | 169.8202 | [8 empty] | 1.4548 | 184.791 | 1.6134 | 159.9371 | 1.8409 | 113.4696
|  | 1.5247 | 169.1463 | [8 empty] | 1.4549 | 184.7415 | 1.6167 | 159.4261 | 1.8413 | 113.3152
|  | 1.5279 | 168.4722 | [8 empty] | 1.455 | 184.692 | 1.6199 | 158.9147 | 1.8418 | 113.161
|  | 1.5312 | 167.7979 | [8 empty] | 1.455 | 184.6425 | 1.6231 | 158.4029 | 1.8422 | 113.0071
|  | 1.5344 | 167.1234 | [8 empty] | 1.4551 | 184.5931 | 1.6264 | 157.8907 | 1.8426 | 112.8534
|  | 1.5377 | 166.4488 | [8 empty] | 1.4552 | 184.5438 | 1.6296 | 157.3782 | 1.8431 | 112.6999
|  | 1.5409 | 165.774 | [8 empty] | 1.4552 | 184.4945 | 1.6329 | 156.8652 | 1.8435 | 112.5467
|  | 1.5441 | 165.0989 | [8 empty] | 1.4553 | 184.4453 | 1.6361 | 156.3518 | 1.844 | 112.3937
|  | 1.5474 | 164.4238 | [8 empty] | 1.4554 | 184.3961 | 1.6393 | 155.8381 | 1.8444 | 112.241
|  | 1.5506 | 163.7484 | [8 empty] | 1.4554 | 184.347 | 1.6426 | 155.324 | 1.8448 | 112.0885
|  | 1.5539 | 163.0729 | [8 empty] | 1.4555 | 184.2979 | 1.6458 | 154.8095 | 1.8453 | 111.9362
|  | 1.5571 | 162.3973 | [8 empty] | 1.4556 | 184.2489 | 1.6491 | 154.2947 | 1.8457 | 111.7842
|  | 1.5603 | 161.7215 | [8 empty] | 1.4556 | 184.2 | 1.6523 | 153.7795 | 1.8461 | 111.6324
|  | 1.5636 | 161.0455 | [8 empty] | 1.4557 | 184.1511 | 1.6555 | 153.2639 | 1.8461 | 111.6324
|  | 1.5668 | 160.3694 | [8 empty] | 1.4558 | 184.1023 | 1.6588 | 152.748 | 1.8462 | 111.6093
|  | 1.5701 | 159.6932 | [8 empty] | 1.4558 | 184.0535 | 1.662 | 152.2317 | 1.8463 | 111.587
|  | 1.5733 | 159.0169 | [8 empty] | 1.4559 | 184.0048 | 1.6653 | 151.7151 | 1.8464 | 111.565
|  | 1.5765 | 158.3404 | [8 empty] | 1.456 | 183.9561 | 1.6685 | 151.1982 | 1.8465 | 111.5424
|  | 1.5798 | 157.6638 | [8 empty] | 1.456 | 183.9075 | 1.6717 | 150.6809 | 1.8466 | 111.5202
|  | 1.583 | 156.987 | [8 empty] | 1.4561 | 183.8589 | 1.675 | 150.1634 | 1.8467 | 111.4972
|  | 1.5863 | 156.3102 | [8 empty] | 1.4562 | 183.8104 | 1.6782 | 149.6454 | 1.8467 | 111.475
|  | 1.5895 | 155.6332 | [8 empty] | 1.4562 | 183.762 | 1.6815 | 149.1272 | 1.8468 | 111.4521
|  | 1.5927 | 154.9562 | [8 empty] | 1.4563 | 183.7136 | 1.6847 | 148.6087 | 1.8469 | 111.4298
|  | 1.596 | 154.279 | [8 empty] | 1.4564 | 183.6653 | 1.6879 | 148.0898 | 1.847 | 111.407
|  | 1.5992 | 153.6018 | [8 empty] | 1.4564 | 183.617 | 1.6912 | 147.5707 | 1.8471 | 111.3845
|  | 1.6025 | 152.9244 | [8 empty] | 1.4565 | 183.5687 | 1.6944 | 147.0512 | 1.8472 | 111.3618
|  | 1.6057 | 152.247 | [8 empty] | 1.4566 | 183.5206 | 1.6977 | 146.5315 | 1.8473 | 111.3393
|  | 1.6089 | 151.5695 | [8 empty] | 1.4566 | 183.4725 | 1.7009 | 146.0114 | 1.8474 | 111.3166
|  | 1.6122 | 150.8919 | [8 empty] | 1.4567 | 183.4244 | 1.7042 | 145.4911 | 1.8474 | 111.294
|  | 1.6154 | 150.2143 | [8 empty] | 1.4568 | 183.3764 | 1.7074 | 144.9705 | 1.8475 | 111.2713
|  | 1.6187 | 149.5365 | [8 empty] | 1.4568 | 183.3284 | 1.7106 | 144.4496 | 1.8476 | 111.2487
|  | 1.6219 | 148.8587 | [8 empty] | 1.4569 | 183.2805 | 1.7139 | 143.9285 | 1.8477 | 111.2261
|  | 1.6251 | 148.1809 | [8 empty] | 1.4602 | 182.5641 | 1.7171 | 143.4071 | 1.8478 | 111.2034
|  | 1.6284 | 147.503 | [8 empty] | 1.4634 | 181.8475 | 1.7204 | 142.8854 | 1.8479 | 111.1807
|  | 1.6316 | 146.825 | [8 empty] | 1.4666 | 181.1307 | 1.7236 | 142.3634 | 1.848 | 111.1581
|  | 1.6349 | 146.147 | [8 empty] | 1.4699 | 180.4138 | 1.7268 | 141.8412 | 1.8481 | 111.1354
|  | 1.6381 | 145.469 | [8 empty] | 1.4731 | 179.6968 | 1.7301 | 141.3188 | 1.8481 | 111.1127
|  | 1.6414 | 144.7909 | [8 empty] | 1.4764 | 178.9796 | 1.7333 | 140.7961 | 1.8482 | 111.09
|  | 1.6446 | 144.1128 | [8 empty] | 1.4796 | 178.2623 | 1.7366 | 140.2732 | 1.8483 | 111.0673
|  | 1.6478 | 143.4347 | [8 empty] | 1.4828 | 177.5449 | 1.7398 | 139.75 | 1.8484 | 111.0446
|  | 1.6511 | 142.7566 | [8 empty] | 1.4861 | 176.8273 | 1.743 | 139.2266 | 1.8485 | 111.0219
|  | 1.6543 | 142.0784 | [8 empty] | 1.4893 | 176.1096 | 1.7463 | 138.703 | 1.8486 | 110.9992
|  | 1.6576 | 141.4002 | [8 empty] | 1.4926 | 175.3918 | 1.7495 | 138.1792 | 1.8487 | 110.9765
|  | 1.6608 | 140.7221 | [8 empty] | 1.4958 | 174.6739 | 1.7528 | 137.6551 | 1.8488 | 110.9538
|  | 1.664 | 140.0439 | [8 empty] | 1.499 | 173.9559 | 1.756 | 137.1308 | 1.8488 | 110.9311
|  | 1.6673 | 139.3657 | [8 empty] | 1.5023 | 173.2377 | 1.7592 | 136.6064 | 1.8489 | 110.9083
|  | 1.6705 | 138.6875 | [8 empty] | 1.5055 | 172.5195 | 1.7625 | 136.0817 | 1.849 | 110.8856
|  | 1.6738 | 138.0093 | [8 empty] | 1.5088 | 171.8011 | 1.7657 | 135.5568 | 1.8491 | 110.8629
|  | 1.677 | 137.3312 | [8 empty] | 1.512 | 171.0827 | 1.769 | 135.0317 | 1.8492 | 110.8401
|  | 1.6802 | 136.6531 | [8 empty] | 1.5152 | 170.3642 | 1.7722 | 134.5064 | 1.8493 | 110.8174
|  | 1.6835 | 135.975 | [8 empty] | 1.5185 | 169.6455 | 1.7754 | 133.981 | 1.8494 | 110.7946
|  | 1.6867 | 135.2969 | [8 empty] | 1.5217 | 168.9268 | 1.7787 | 133.4553 | 1.8495 | 110.7718
|  | 1.69 | 134.6189 | [8 empty] | 1.525 | 168.208 | 1.7819 | 132.9295 | 1.8495 | 110.7491
|  | 1.6932 | 133.9409 | [8 empty] | 1.5282 | 167.4891 | 1.7852 | 132.4035 | 1.8496 | 110.7263
|  | 1.6964 | 133.2629 | [8 empty] | 1.5314 | 166.7702 | 1.7884 | 131.8773 | 1.8497 | 110.7035
|  | 1.6997 | 132.585 | [8 empty] | 1.5347 | 166.0512 | 1.7916 | 131.351 | 1.8498 | 110.6807
|  | 1.7029 | 131.9072 | [8 empty] | 1.5379 | 165.3321 | 1.7949 | 130.8245 | 1.8499 | 110.6579
|  | 1.7062 | 131.2294 | [8 empty] | 1.5412 | 164.6129 | 1.7981 | 130.2979 | 1.85 | 110.6351
|  | 1.7094 | 130.5517 | [8 empty] | 1.5444 | 163.8937 | 1.8014 | 129.771 | 1.8501 | 110.6123
|  | 1.7126 | 129.8741 | [8 empty] | 1.5476 | 163.1744 | 1.8046 | 129.2441 | 1.8501 | 110.5895
|  | 1.7159 | 129.1965 | [8 empty] | 1.5509 | 162.4551 | 1.8078 | 128.717 | 1.8502 | 110.5667
|  | 1.7191 | 128.519 | [8 empty] | 1.5541 | 161.7358 | 1.8111 | 128.1898 | 1.8503 | 110.5439
|  | 1.7224 | 127.8416 | [8 empty] | 1.5574 | 161.0164 | 1.8143 | 127.6624 | 1.8504 | 110.5211
|  | 1.7256 | 127.1642 | [8 empty] | 1.5606 | 160.2969 | 1.8176 | 127.1349 | 1.8505 | 110.4983
|  | 1.7288 | 126.487 | [8 empty] | 1.5638 | 159.5774 | 1.8208 | 126.6072 | 1.8506 | 110.4754
|  | 1.7321 | 125.8098 | [8 empty] | 1.5671 | 158.8579 | 1.824 | 126.0795 | 1.8507 | 110.4526
|  | 1.7353 | 125.1328 | [8 empty] | 1.5703 | 158.1384 | 1.8273 | 125.5516 | 1.8508 | 110.4298
|  | 1.7386 | 124.4559 | [8 empty] | 1.5736 | 157.4188 | 1.8305 | 125.0236 | 1.8508 | 110.4069
|  | 1.7418 | 123.779 | [8 empty] | 1.5768 | 156.6993 | 1.8338 | 124.4955 | 1.8509 | 110.3841
|  | 1.745 | 123.1023 | [8 empty] | 1.58 | 155.9797 | 1.837 | 123.9673 | 1.851 | 110.3612
|  | 1.7483 | 122.4257 | [8 empty] | 1.5833 | 155.2601 | 1.8402 | 123.439 | 1.8511 | 110.3383
|  | 1.7515 | 121.7493 | [8 empty] | 1.5865 | 154.5405 | 1.8435 | 122.9106 | 1.8512 | 110.3155
|  | 1.7548 | 121.0729 | [8 empty] | 1.5898 | 153.8209 | 1.8467 | 122.3821 | 1.8513 | 110.2926
|  | 1.758 | 120.3967 | [8 empty] | 1.593 | 153.1013 | 1.85 | 121.8535 | 1.8514 | 110.2697
|  | 1.7612 | 119.7207 | [8 empty] | 1.5962 | 152.3817 | 1.8532 | 121.3248 | 1.8515 | 110.2469
|  | 1.7645 | 119.0448 | [8 empty] | 1.5995 | 151.6622 | 1.8564 | 120.796 | 1.8515 | 110.224
|  | 1.7677 | 118.369 | [8 empty] | 1.6027 | 150.9426 | 1.8597 | 120.2672 | 1.8516 | 110.2011
|  | 1.771 | 117.6934 | [8 empty] | 1.606 | 150.2231 | 1.8629 | 119.7382 | 1.8517 | 110.1782
|  | 1.7742 | 117.0179 | [8 empty] | 1.6092 | 149.5036 | 1.8662 | 119.2093 | 1.8518 | 110.1553
|  | 1.7774 | 116.3426 | [8 empty] | 1.6124 | 148.7841 | 1.8694 | 118.6802 | 1.8519 | 110.1324
|  | 1.7807 | 115.6675 | [8 empty] | 1.6157 | 148.0647 | 1.8726 | 118.1511 | 1.852 | 110.1095
|  | 1.7839 | 114.9925 | [8 empty] | 1.6189 | 147.3453 | 1.8759 | 117.6219 | 1.8521 | 110.0866
|  | 1.7872 | 114.3177 | [8 empty] | 1.6222 | 146.6259 | 1.8791 | 117.0927 | 1.8522 | 110.0637
|  | 1.7904 | 113.6431 | [8 empty] | 1.6254 | 145.9066 | 1.8824 | 116.5634 | 1.8522 | 110.0407
|  | 1.7936 | 112.9687 | [8 empty] | 1.6286 | 145.1874 | 1.8856 | 116.0341 | 1.8523 | 110.0178
|  | 1.7969 | 112.2945 | [8 empty] | 1.6319 | 144.4682 | 1.8889 | 115.5047 | 1.8524 | 109.9949
|  | 1.8001 | 111.6204 | [8 empty] | 1.6351 | 143.7491 | 1.8921 | 114.9753 | 1.8525 | 109.9719
|  | 1.8034 | 110.9466 | [8 empty] | 1.6384 | 143.03 | 1.8953 | 114.4459 | 1.8526 | 109.949
|  | 1.8066 | 110.2729 | [8 empty] | 1.6416 | 142.311 | 1.8986 | 113.9164 | 1.8527 | 109.9261
|  | 1.8099 | 109.5995 | [8 empty] | 1.6449 | 141.5921 | 1.9018 | 113.3869 | 1.8528 | 109.9031
|  | 1.8131 | 108.9263 | [8 empty] | 1.6481 | 140.8733 | 1.9051 | 112.8574 | 1.8529 | 109.8802
|  | 1.8163 | 108.2533 | [8 empty] | 1.6513 | 140.1545 | 1.9083 | 112.3278 | 1.8529 | 109.8572
|  | 1.8196 | 107.5805 | [8 empty] | 1.6546 | 139.4358 | 1.9115 | 111.7983 | 1.853 | 109.8343
|  | 1.8228 | 106.908 | [8 empty] | 1.6578 | 138.7173 | 1.9148 | 111.2687 | 1.8531 | 109.8113
|  | 1.8261 | 106.2356 | [8 empty] | 1.6611 | 137.9988 | 1.918 | 110.7391 | 1.8532 | 109.7883
|  | 1.8293 | 105.5636 | [8 empty] | 1.6643 | 137.2804 | 1.9213 | 110.2096 | 1.8533 | 109.7654
|  | 1.8325 | 104.8917 | [8 empty] | 1.6675 | 136.5622 | 1.9245 | 109.68 | 1.8534 | 109.7424
|  | 1.8358 | 104.2201 | [8 empty] | 1.6708 | 135.844 | 1.9277 | 109.1504 | 1.8535 | 109.7194
|  | 1.839 | 103.5488 | [8 empty] | 1.674 | 135.126 | 1.931 | 108.6209 | 1.8536 | 109.6964
|  | 1.8423 | 102.8777 | [8 empty] | 1.6773 | 134.4081 | 1.9342 | 108.0913 | 1.8536 | 109.6734
|  | 1.8455 | 102.2068 | [8 empty] | 1.6805 | 133.6903 | 1.9375 | 107.5618 | 1.8537 | 109.6504
|  | 1.8487 | 101.5363 | [8 empty] | 1.6837 | 132.9727 | 1.9407 | 107.0323 | 1.8538 | 109.6274
|  | 1.852 | 100.8659 | [8 empty] | 1.687 | 132.2552 | 1.9439 | 106.5028 | 1.8539 | 109.6044
|  | 1.8552 | 100.1959 | [8 empty] | 1.6902 | 131.5378 | 1.9472 | 105.9733 | 1.854 | 109.5814
|  | 1.8585 | 99.5261 | [8 empty] | 1.6935 | 130.8206 | 1.9504 | 105.4439 | 1.8541 | 109.5584
|  | 1.8617 | 98.8567 | [8 empty] | 1.6967 | 130.1035 | 1.9537 | 104.9145 | 1.8542 | 109.5354
|  | 1.8649 | 98.1875 | [8 empty] | 1.6999 | 129.3866 | 1.9569 | 104.3851 | 1.8543 | 109.5124
|  | 1.8682 | 97.5185 | [8 empty] | 1.7032 | 128.6698 | 1.9601 | 103.8558 | 1.8543 | 109.4894
|  | 1.8714 | 96.8499 | [8 empty] | 1.7064 | 127.9532 | 1.9634 | 103.3265 | 1.8544 | 109.4664
|  | 1.8747 | 96.1816 | [8 empty] | 1.7097 | 127.2368 | 1.9666 | 102.7973 | 1.8545 | 109.4433
|  | 1.8779 | 95.5136 | [8 empty] | 1.7129 | 126.5205 | 1.9699 | 102.2681 | 1.8546 | 109.4203
|  | 1.8811 | 94.8459 | [8 empty] | 1.7161 | 125.8044 | 1.9731 | 101.739 | 1.8547 | 109.3973
|  | 1.8844 | 94.1785 | [8 empty] | 1.7194 | 125.0885 | 1.9763 | 101.21 | 1.8548 | 109.3742
|  | 1.8876 | 93.5114 | [8 empty] | 1.7226 | 124.3728 | 1.9796 | 100.681 | 1.8549 | 109.3512
|  | 1.8909 | 92.8447 | [8 empty] | 1.7259 | 123.6573 | 1.9828 | 100.1521 | 1.855 | 109.3282
|  | 1.8941 | 92.1782 | [8 empty] | 1.7291 | 122.9419 | 1.9861 | 99.6232 | 1.855 | 109.3051
|  | 1.8973 | 91.5121 | [8 empty] | 1.7323 | 122.2268 | 1.9893 | 99.0945 | 1.8551 | 109.2821
|  | 1.9006 | 90.8464 | [8 empty] | 1.7356 | 121.5119 | 1.9925 | 98.5658 | 1.8552 | 109.259
|  | 1.9038 | 90.1809 | [8 empty] | 1.7388 | 120.7972 | 1.9958 | 98.0372 | 1.8553 | 109.2359
|  | 1.9071 | 89.5159 | [8 empty] | 1.7421 | 120.0827 | 1.999 | 97.5087 | 1.8554 | 109.2129
|  | 1.9103 | 88.8511 | [8 empty] | 1.7453 | 119.3684 | 2.0023 | 96.9802 | 1.8555 | 109.1908
|  | 1.9135 | 88.1867 | [8 empty] | 1.7485 | 118.6543 | 2.0055 | 96.4519 | 1.8556 | 109.1681
|  | 1.9168 | 87.5227 | [8 empty] | 1.7518 | 117.9405 | 2.0087 | 95.9237 | 1.8557 | 109.1458
|  | 1.92 | 86.859 | [8 empty] | 1.755 | 117.2269 | 2.012 | 95.3955 | 1.8557 | 109.1232
|  | 1.9233 | 86.1957 | [8 empty] | 1.7583 | 116.5135 | 2.0152 | 94.8675 | 1.8558 | 109.1008
|  | 1.9265 | 85.5328 | [8 empty] | 1.7615 | 115.8004 | 2.0185 | 94.3396 | 1.8559 | 109.0782
|  | 1.9297 | 84.8702 | [8 empty] | 1.7647 | 115.0876 | 2.0217 | 93.8118 | 1.856 | 109.0558
|  | 1.933 | 84.208 | [8 empty] | 1.768 | 114.3749 | 2.0249 | 93.2841 | 1.8561 | 109.0333
|  | 1.9362 | 83.5462 | [8 empty] | 1.7712 | 113.6626 | 2.0282 | 92.7565 | 1.8562 | 109.0108
|  | 1.9395 | 82.8848 | [8 empty] | 1.7745 | 112.9505 | 2.0314 | 92.229 | 1.8563 | 108.9883
|  | 1.9427 | 82.2238 | [8 empty] | 1.7777 | 112.2386 | 2.0347 | 91.7017 | 1.8563 | 108.9658
|  | 1.9459 | 81.5631 | [8 empty] | 1.7809 | 111.5271 | 2.0379 | 91.1745 | 1.8564 | 108.9433
|  | 1.9492 | 80.9029 | [8 empty] | 1.7842 | 110.8158 | 2.0411 | 90.6474 | 1.8565 | 108.9208
|  | 1.9524 | 80.2431 | [8 empty] | 1.7874 | 110.1048 | 2.0444 | 90.1205 | 1.8566 | 108.8983
|  | 1.9557 | 79.5836 | [8 empty] | 1.7907 | 109.394 | 2.0476 | 89.5937 | 1.8567 | 108.8758
|  | 1.9589 | 78.9246 | [8 empty] | 1.7939 | 108.6836 | 2.0509 | 89.0671 | 1.8568 | 108.8533
|  | 1.9621 | 78.266 | [8 empty] | 1.7971 | 107.9734 | 2.0541 | 88.5406 | 1.8569 | 108.8308
|  | 1.9654 | 77.6079 | [8 empty] | 1.8004 | 107.2636 | 2.0574 | 88.0142 | 1.857 | 108.8082
|  | 1.9686 | 76.9501 | [8 empty] | 1.8036 | 106.5541 | 2.0606 | 87.488 | 1.857 | 108.7857
|  | 1.9719 | 76.2928 | [8 empty] | 1.8069 | 105.8448 | 2.0638 | 86.962 | 1.8571 | 108.7632
|  | 1.9751 | 75.6359 | [8 empty] | 1.8101 | 105.1359 | 2.0671 | 86.4361 | 1.8572 | 108.7407
|  | 1.9783 | 74.9795 | [8 empty] | 1.8134 | 104.4273 | 2.0703 | 85.9104 | 1.8573 | 108.7181
|  | 1.9816 | 74.3235 | [8 empty] | 1.8166 | 103.719 | 2.0736 | 85.3849 | 1.8574 | 108.6956
|  | 1.9848 | 73.6679 | [8 empty] | 1.8198 | 103.011 | 2.0768 | 84.8595 | 1.8575 | 108.6731
|  | 1.9881 | 73.0128 | [8 empty] | 1.8231 | 102.3034 | 2.08 | 84.3343 | 1.8576 | 108.6505
|  | 1.9913 | 72.3581 | [8 empty] | 1.8263 | 101.5961 | 2.0833 | 83.8093 | 1.8577 | 108.628
|  | 1.9946 | 71.7039 | [8 empty] | 1.8296 | 100.8891 | 2.0865 | 83.2844 | 1.8577 | 108.6054
|  | 1.9978 | 71.0502 | [8 empty] | 1.8328 | 100.1825 | 2.0898 | 82.7598 | 1.8578 | 108.5829
|  | 2.001 | 70.3969 | [8 empty] | 1.836 | 99.4762 | 2.093 | 82.2353 | 1.8579 | 108.5604
|  | 2.0043 | 69.7441 | [8 empty] | 1.8393 | 98.7703 | 2.0962 | 81.711 | 1.858 | 108.5378
|  | 2.0075 | 69.0918 | [8 empty] | 1.8425 | 98.0647 | 2.0995 | 81.1869 | 1.8581 | 108.5153
|  | 2.0108 | 68.44 | [8 empty] | 1.8458 | 97.3595 | 2.1027 | 80.663 | 1.8582 | 108.4927
|  | 2.014 | 67.7886 | [8 empty] | 1.849 | 96.6547 | 2.106 | 80.1393 | 1.8583 | 108.4701
|  | 2.0172 | 67.1377 | [8 empty] | 1.8522 | 95.9502 | 2.1092 | 79.6159 | 1.8584 | 108.4476
|  | 2.0205 | 66.4874 | [8 empty] | 1.8555 | 95.2462 | 2.1124 | 79.0926 | 1.8584 | 108.425
|  | 2.0237 | 65.8375 | [8 empty] | 1.8587 | 94.5424 | 2.1157 | 78.5695 | 1.8585 | 108.4024
|  | 2.027 | 65.1881 | [8 empty] | 1.862 | 93.8391 | 2.1189 | 78.0467 | 1.8586 | 108.3799
|  | 2.0302 | 64.5392 | [8 empty] | 1.8652 | 93.1362 | 2.1222 | 77.524 | 1.8587 | 108.3573
|  | 2.0334 | 63.8908 | [8 empty] | 1.8684 | 92.4337 | 2.1254 | 77.0016 | 1.8588 | 108.3347
|  | 2.0367 | 63.243 | [8 empty] | 1.8717 | 91.7315 | 2.1286 | 76.4794 | 1.8589 | 108.3122
|  | 2.0399 | 62.5956 | [8 empty] | 1.8749 | 91.0298 | 2.1319 | 75.9575 | 1.859 | 108.2896
|  | 2.0432 | 61.9488 | [8 empty] | 1.8782 | 90.3285 | 2.1351 | 75.4358 | 1.8591 | 108.267
|  | 2.0464 | 61.3025 | [8 empty] | 1.8814 | 89.6275 | 2.1384 | 74.9143 | 1.8591 | 108.2444
|  | 2.0496 | 60.6568 | [8 empty] | 1.8846 | 88.927 | 2.1416 | 74.393 | 1.8592 | 108.2218
|  | 2.0529 | 60.0115 | [8 empty] | 1.8879 | 88.227 | 2.1448 | 73.872 | 1.8593 | 108.1993
|  | 2.0561 | 59.3668 | [8 empty] | 1.8911 | 87.5273 | 2.1481 | 73.3512 | 1.8594 | 108.1767
|  | 2.0594 | 58.7227 | [8 empty] | 1.8944 | 86.8281 | 2.1513 | 72.8307 | 1.8595 | 108.1541
|  | 2.0626 | 58.0791 | [8 empty] | 1.8976 | 86.1293 | 2.1546 | 72.3104 | 1.8596 | 108.1315
|  | 2.0658 | 57.436 | [8 empty] | 1.9008 | 85.4309 | 2.1578 | 71.7904 | 1.8597 | 108.1089
|  | 2.0691 | 56.7935 | [8 empty] | 1.9041 | 84.733 | 2.161 | 71.2706 | 1.8598 | 108.0863
|  | 2.0723 | 56.1515 | [8 empty] | 1.9073 | 84.0356 | 2.1643 | 70.7511 | 1.8598 | 108.0637
|  | 2.0756 | 55.5101 | [8 empty] | 1.9106 | 83.3386 | 2.1675 | 70.2319 | 1.8599 | 108.0411
|  | 2.0788 | 54.8693 | [8 empty] | 1.9138 | 82.642 | 2.1708 | 69.7129 | 1.86 | 108.0185
|  | 2.082 | 54.2291 | [8 empty] | 1.917 | 81.9459 | 2.174 | 69.1942 | 1.8601 | 107.9959
|  | 2.0853 | 53.5894 | [8 empty] | 1.9203 | 81.2503 | 2.1772 | 68.6758 | 1.8602 | 107.9733
|  | 2.0885 | 52.9503 | [8 empty] | 1.9235 | 80.5551 | 2.1805 | 68.1576 | 1.8603 | 107.9507
|  | 2.0918 | 52.3117 | [8 empty] | 1.9268 | 79.8605 | 2.1837 | 67.6397 | 1.8604 | 107.928
|  | 2.095 | 51.6738 | [8 empty] | 1.93 | 79.1662 | 2.187 | 67.1221 | 1.8605 | 107.9054
|  | 2.0982 | 51.0364 | [8 empty] | 1.9332 | 78.4725 | 2.1902 | 66.6048 | 1.8605 | 107.8828
|  | 2.1015 | 50.3997 | [8 empty] | 1.9365 | 77.7793 | 2.1934 | 66.0878 | 1.8606 | 107.8602
|  | 2.1047 | 49.7635 | [8 empty] | 1.9397 | 77.0866 | 2.1967 | 65.5711 | 1.8607 | 107.8376
|  | 2.108 | 49.1279 | [8 empty] | 1.943 | 76.3943 | 2.1999 | 65.0546 | 1.8608 | 107.8149
|  | 2.1112 | 48.493 | [8 empty] | 1.9462 | 75.7026 | 2.2032 | 64.5385 | 1.8609 | 107.7923
|  | 2.1144 | 47.8586 | [8 empty] | 1.9494 | 75.0113 | 2.2064 | 64.0226 | 1.861 | 107.7697
|  | 2.1177 | 47.2249 | [8 empty] | 1.9527 | 74.3206 | 2.2096 | 63.5071 | 1.8611 | 107.7471
|  | 2.1209 | 46.5918 | [8 empty] | 1.9559 | 73.6304 | 2.2129 | 62.9918 | 1.8612 | 107.7244
|  | 2.1242 | 45.9592 | [8 empty] | 1.9592 | 72.9407 | 2.2161 | 62.4769 | 1.8612 | 107.7018
|  | 2.1274 | 45.3274 | [8 empty] | 1.9624 | 72.2515 | 2.2194 | 61.9623 | 1.8613 | 107.6792
|  | 2.1306 | 44.6961 | [8 empty] | 1.9656 | 71.5629 | 2.2226 | 61.448 | 1.8614 | 107.6565
|  | 2.1339 | 44.0655 | [8 empty] | 1.9689 | 70.8748 | 2.2258 | 60.934 | 1.8615 | 107.6339
|  | 2.1371 | 43.4355 | [8 empty] | 1.9721 | 70.1872 | 2.2291 | 60.4204 | 1.8616 | 107.6112
|  | 2.1404 | 42.8061 | [8 empty] | 1.9754 | 69.5002 | 2.2323 | 59.907 | 1.8617 | 107.5886
|  | 2.1436 | 42.1774 | [8 empty] | 1.9786 | 68.8137 | 2.2356 | 59.394 | 1.8618 | 107.566
|  | 2.1468 | 41.5494 | [8 empty] | 1.9818 | 68.1277 | 2.2388 | 58.8814 | 1.8619 | 107.5433
|  | 2.1501 | 40.922 | [8 empty] | 1.9851 | 67.4423 | 2.2421 | 58.369 | 1.8619 | 107.5207
|  | 2.1533 | 40.2952 | [8 empty] | 1.9883 | 66.7575 | 2.2453 | 57.857 | 1.862 | 107.498
|  | 2.1566 | 39.6691 | [8 empty] | 1.9916 | 66.0732 | 2.2485 | 57.3453 | 1.8621 | 107.4754
|  | 2.1598 | 39.0437 | [8 empty] | 1.9948 | 65.3895 | 2.2518 | 56.834 | 1.8622 | 107.4527
|  | 2.1631 | 38.4189 | [8 empty] | 1.9981 | 64.7064 | 2.255 | 56.323 | 1.8623 | 107.4301
|  | 2.1663 | 37.7948 | [8 empty] | 2.0013 | 64.0238 | 2.2583 | 55.8124 | 1.8624 | 107.4074
|  | 2.1695 | 37.1714 | [8 empty] | 2.0045 | 63.3418 | 2.2615 | 55.3021 | 1.8625 | 107.3847
|  | 2.1728 | 36.5486 | [8 empty] | 2.0078 | 62.6604 | 2.2647 | 54.7922 | 1.8625 | 107.3621
|  | 2.176 | 35.9265 | [8 empty] | 2.011 | 61.9796 | 2.268 | 54.2826 | 1.8626 | 107.3394
|  | 2.1793 | 35.3051 | [8 empty] | 2.0143 | 61.2994 | 2.2712 | 53.7734 | 1.8627 | 107.3168
|  | 2.1825 | 34.6844 | [8 empty] | 2.0175 | 60.6198 |  |  | 1.8628 | 107.2941
|  | 2.1857 | 34.0644 | [8 empty] | 2.0207 | 59.9407 |  |  | 1.8629 | 107.2714
|  | 2.189 | 33.4451 | [8 empty] | 2.024 | 59.2623 |  |  | 1.863 | 107.2487
|  | 2.1922 | 32.8265 | [8 empty] | 2.0272 | 58.5845 |  |  | 1.8631 | 107.2261
|  | 2.1955 | 32.2086 | [8 empty] | 2.0305 | 57.9073 |  |  | 1.8632 | 107.2034
|  | 2.1987 | 31.5914 | [8 empty] | 2.0337 | 57.2307 |  |  | 1.8632 | 107.1807
|  | 2.2019 | 30.9749 | [8 empty] | 2.0369 | 56.5547 |  |  | 1.8633 | 107.1581
|  | 2.2052 | 30.3591 | [8 empty] | 2.0402 | 55.8794 |  |  | 1.8634 | 107.1354
|  | 2.2084 | 29.744 | [8 empty] | 2.0434 | 55.2047 |  |  | 1.8635 | 107.1127
|  | 2.2117 | 29.1297 | [8 empty] | 2.0467 | 54.5306 |  |  | 1.8636 | 107.09
|  | 2.2149 | 28.516 | [8 empty] | 2.0499 | 53.8572 |  |  | 1.8637 | 107.0673
|  | 2.2181 | 27.9031 | [8 empty] | 2.0531 | 53.1844 |  |  | 1.8638 | 107.0447
|  | 2.2214 | 27.291 | [8 empty] | 2.0564 | 52.5122 |  |  | 1.8639 | 107.022
|  | 2.2246 | 26.6796 | [8 empty] | 2.0596 | 51.8407 |  |  | 1.8639 | 106.9993
|  | 2.2279 | 26.0689 | [8 empty] | 2.0629 | 51.1699 |  |  | 1.864 | 106.9766
|  | 2.2311 | 25.4589 | [8 empty] | 2.0661 | 50.4997 |  |  | 1.8641 | 106.9539
|  | 2.2343 | 24.8497 | [8 empty] | 2.0693 | 49.8301 |  |  | 1.8642 | 106.9312
|  | 2.2376 | 24.2413 | [8 empty] | 2.0726 | 49.1613 |  |  | 1.8643 | 106.9085
|  | 2.2408 | 23.6336 | [8 empty] | 2.0758 | 48.4931 |  |  | 1.8644 | 106.8858
|  | 2.2441 | 23.0266 | [8 empty] | 2.0791 | 47.8256 |  |  | 1.8645 | 106.8631
|  | 2.2473 | 22.4204 | [8 empty] | 2.0823 | 47.1587 |  |  | 1.8646 | 106.8404
|  | 2.2505 | 21.815 | [8 empty] | 2.0855 | 46.4926 |  |  | 1.8646 | 106.8177
|  | 2.2538 | 21.2103 | [8 empty] | 2.0888 | 45.8271 |  |  | 1.8647 | 106.795
|  | 2.257 | 20.6065 | [8 empty] | 2.092 | 45.1623 |  |  | 1.8648 | 106.7723
|  | 2.2603 | 20.0034 | [8 empty] | 2.0953 | 44.4982 |  |  | 1.8649 | 106.7496
|  | 2.2635 | 19.401 | [8 empty] | 2.0985 | 43.8349 |  |  | 1.865 | 106.7269
|  | 2.2667 | 18.7995 | [8 empty] | 2.1017 | 43.1722 |  |  | 1.8651 | 106.7042
|  | 2.27 | 18.1987 | [8 empty] | 2.105 | 42.5102 |  |  | 1.8652 | 106.6815
[12 empty] | 2.1082 | 41.8489 |  |  | 1.8653 | 106.6588
[12 empty] | 2.1115 | 41.1884 |  |  | 1.8653 | 106.6361
[12 empty] | 2.1147 | 40.5285 |  |  | 1.8654 | 106.6134
[12 empty] | 2.1179 | 39.8694 |  |  | 1.8655 | 106.5907
[12 empty] | 2.1212 | 39.2111 |  |  | 1.8656 | 106.568
[12 empty] | 2.1244 | 38.5534 |  |  | 1.8657 | 106.5452
[12 empty] | 2.1277 | 37.8965 |  |  | 1.8658 | 106.5225
[12 empty] | 2.1309 | 37.2403 |  |  | 1.8659 | 106.4998
[12 empty] | 2.1341 | 36.5849 |  |  | 1.866 | 106.4771
[12 empty] | 2.1374 | 35.9302 |  |  | 1.866 | 106.4544
[12 empty] | 2.1406 | 35.2762 |  |  | 1.8661 | 106.4317
[12 empty] | 2.1439 | 34.6231 |  |  | 1.8662 | 106.4089
[12 empty] | 2.1471 | 33.9706 |  |  | 1.8663 | 106.3862
[12 empty] | 2.1503 | 33.319 |  |  | 1.8664 | 106.3635
[12 empty] | 2.1536 | 32.6681 |  |  | 1.8665 | 106.3408
[12 empty] | 2.1568 | 32.0179 |  |  | 1.8666 | 106.318
[12 empty] | 2.1601 | 31.3686 |  |  | 1.8667 | 106.2953
[12 empty] | 2.1633 | 30.72 |  |  | 1.8667 | 106.2726
[12 empty] | 2.1666 | 30.0722 |  |  | 1.8668 | 106.2498
[12 empty] | 2.1698 | 29.4252 |  |  | 1.8669 | 106.2271
[12 empty] | 2.173 | 28.779 |  |  | 1.867 | 106.2044
[12 empty] | 2.1763 | 28.1335 |  |  | 1.8671 | 106.1817
[12 empty] | 2.1795 | 27.4889 |  |  | 1.8672 | 106.1589
[12 empty] | 2.1828 | 26.8451 |  |  | 1.8673 | 106.1362
[12 empty] | 2.186 | 26.202 |  |  | 1.8674 | 106.1134
[12 empty] | 2.1892 | 25.5598 |  |  | 1.8674 | 106.0907
[12 empty] | 2.1925 | 24.9184 |  |  | 1.8675 | 106.068
[12 empty] | 2.1957 | 24.2778 |  |  | 1.8676 | 106.0452
[12 empty] | 2.199 | 23.638 |  |  | 1.8677 | 106.0225
[12 empty] | 2.2022 | 22.9991 |  |  | 1.8678 | 105.9997
[12 empty] | 2.2054 | 22.361 |  |  | 1.8679 | 105.977
[12 empty] | 2.2087 | 21.7237 |  |  | 1.868 | 105.9543
[12 empty] | 2.2119 | 21.0872 |  |  | 1.8681 | 105.9315
[12 empty] | 2.2152 | 20.4516 |  |  | 1.8681 | 105.9088
[12 empty] | 2.2184 | 19.8168 |  |  | 1.8682 | 105.886
[12 empty] | 2.2216 | 19.1829 |  |  | 1.8683 | 105.8633
[12 empty] | 2.2249 | 18.5498 |  |  | 1.8684 | 105.8405
[12 empty] | 2.2281 | 17.9176 |  |  | 1.8685 | 105.8178
[12 empty] | 2.2314 | 17.2862 |  |  | 1.8686 | 105.795
[12 empty] | 2.2346 | 16.6557 |  |  | 1.8687 | 105.7723
[12 empty] | 2.2378 | 16.026 |  |  | 1.8687 | 105.7495
[12 empty] | 2.2411 | 15.3972 |  |  | 1.8688 | 105.7268
[12 empty] | 2.2443 | 14.7693 |  |  | 1.8689 | 105.704
[12 empty] | 2.2476 | 14.1423 |  |  | 1.869 | 105.6813
[12 empty] | 2.2508 | 13.5161 |  |  | 1.8691 | 105.6585
[12 empty] | 2.254 | 12.8908 |  |  | 1.8692 | 105.6357
[12 empty] | 2.2573 | 12.2664 |  |  | 1.8693 | 105.613
[12 empty] | 2.2605 | 11.6429 |  |  | 1.8694 | 105.5902
[12 empty] | 2.2638 | 11.0203 |  |  | 1.8694 | 105.5675
[12 empty] | 2.267 | 10.3986 |  |  | 1.8695 | 105.5447
[12 empty] | 2.2702 | 9.7778 |  |  | 1.8696 | 105.522
[16 empty] | 1.8697 | 105.4992
[16 empty] | 1.8698 | 105.4764
[16 empty] | 1.8699 | 105.4537
[16 empty] | 1.87 | 105.4309
[16 empty] | 1.8701 | 105.4081
[16 empty] | 1.8701 | 105.3854
[16 empty] | 1.8702 | 105.3626
[16 empty] | 1.8703 | 105.3398
[16 empty] | 1.8704 | 105.3171
[16 empty] | 1.8705 | 105.2943
[16 empty] | 1.8706 | 105.2715
[16 empty] | 1.8707 | 105.2488
[16 empty] | 1.8708 | 105.226
[16 empty] | 1.8708 | 105.2032
[16 empty] | 1.8709 | 105.1805
[16 empty] | 1.871 | 105.1577
[16 empty] | 1.8711 | 105.1349
[16 empty] | 1.8712 | 105.1121
[16 empty] | 1.8713 | 105.0894
[16 empty] | 1.8714 | 105.0666
[16 empty] | 1.8715 | 105.0438
[16 empty] | 1.8715 | 105.021
[16 empty] | 1.8716 | 104.9983
[16 empty] | 1.8717 | 104.9755
[16 empty] | 1.8718 | 104.9527
[16 empty] | 1.8719 | 104.9299
[16 empty] | 1.872 | 104.9071
[16 empty] | 1.8721 | 104.8844
[16 empty] | 1.8722 | 104.8616
[16 empty] | 1.8722 | 104.8388
[16 empty] | 1.8723 | 104.816
[16 empty] | 1.8724 | 104.7932
[16 empty] | 1.8725 | 104.7705
[16 empty] | 1.8726 | 104.7477
[16 empty] | 1.8727 | 104.7249
[16 empty] | 1.8728 | 104.7021
[16 empty] | 1.8729 | 104.6793
[16 empty] | 1.8729 | 104.6565
[16 empty] | 1.873 | 104.6338
[16 empty] | 1.8731 | 104.611
[16 empty] | 1.8732 | 104.5882
[16 empty] | 1.8733 | 104.5654
[16 empty] | 1.8734 | 104.5426
[16 empty] | 1.8735 | 104.5198
[16 empty] | 1.8736 | 104.497
[16 empty] | 1.8736 | 104.4742
[16 empty] | 1.8737 | 104.4515
[16 empty] | 1.8738 | 104.4287
[16 empty] | 1.8739 | 104.4059
[16 empty] | 1.874 | 104.3831
[16 empty] | 1.8741 | 104.3603
[16 empty] | 1.8742 | 104.3375
[16 empty] | 1.8742 | 104.3147
[16 empty] | 1.8743 | 104.2919
[16 empty] | 1.8744 | 104.2691
[16 empty] | 1.8745 | 104.2463
[16 empty] | 1.8746 | 104.2236
[16 empty] | 1.8747 | 104.2008
[16 empty] | 1.8748 | 104.178
[16 empty] | 1.8749 | 104.1552
[16 empty] | 1.8749 | 104.1324
[16 empty] | 1.875 | 104.1096
[16 empty] | 1.8751 | 104.0868
[16 empty] | 1.8752 | 104.064
[16 empty] | 1.8753 | 104.0412
[16 empty] | 1.8754 | 104.0184
[16 empty] | 1.8755 | 103.9956
[16 empty] | 1.8756 | 103.9728
[16 empty] | 1.8756 | 103.95
[16 empty] | 1.8757 | 103.9272
[16 empty] | 1.8758 | 103.9044
[16 empty] | 1.8759 | 103.8816
[16 empty] | 1.876 | 103.8588
[16 empty] | 1.8761 | 103.836
[16 empty] | 1.8762 | 103.8132
[16 empty] | 1.8763 | 103.7904
[16 empty] | 1.8763 | 103.7676
[16 empty] | 1.8764 | 103.7449
[16 empty] | 1.8765 | 103.7221
[16 empty] | 1.8766 | 103.6993
[16 empty] | 1.8767 | 103.6765
[16 empty] | 1.8768 | 103.6537
[16 empty] | 1.8769 | 103.6309
[16 empty] | 1.877 | 103.6081
[16 empty] | 1.877 | 103.5853
[16 empty] | 1.8771 | 103.5625
[16 empty] | 1.8772 | 103.5397
[16 empty] | 1.8773 | 103.5169
[16 empty] | 1.8774 | 103.4941
[16 empty] | 1.8775 | 103.4713
[16 empty] | 1.8776 | 103.4485
[16 empty] | 1.8777 | 103.4257
[16 empty] | 1.8777 | 103.4029
[16 empty] | 1.8778 | 103.3801
[16 empty] | 1.8779 | 103.3573
[16 empty] | 1.878 | 103.3345
[16 empty] | 1.8781 | 103.3117
[16 empty] | 1.8782 | 103.2889
[16 empty] | 1.8783 | 103.2661
[16 empty] | 1.8784 | 103.2433
[16 empty] | 1.8784 | 103.2205
[16 empty] | 1.8785 | 103.1977
[16 empty] | 1.8786 | 103.1749
[16 empty] | 1.8787 | 103.1521
[16 empty] | 1.8788 | 103.1293
[16 empty] | 1.8789 | 103.1065
[16 empty] | 1.879 | 103.0837
[16 empty] | 1.8791 | 103.0609
[16 empty] | 1.8791 | 103.0381
[16 empty] | 1.8792 | 103.0153
[16 empty] | 1.8793 | 102.9925
[16 empty] | 1.8794 | 102.9697
[16 empty] | 1.8795 | 102.9469
[16 empty] | 1.8796 | 102.9241
[16 empty] | 1.8797 | 102.9013
[16 empty] | 1.8798 | 102.8785
[16 empty] | 1.8798 | 102.8557
[16 empty] | 1.8799 | 102.8329
[16 empty] | 1.88 | 102.8101
[16 empty] | 1.8801 | 102.7873
[16 empty] | 1.8802 | 102.7645
[16 empty] | 1.8803 | 102.7417
[16 empty] | 1.8804 | 102.7189
[16 empty] | 1.8804 | 102.6961
[16 empty] | 1.8805 | 102.6733
[16 empty] | 1.8806 | 102.6505
[16 empty] | 1.8807 | 102.6277
[16 empty] | 1.8808 | 102.6049
[16 empty] | 1.8809 | 102.5821
[16 empty] | 1.881 | 102.5593
[16 empty] | 1.8811 | 102.5365
[16 empty] | 1.8811 | 102.5137
[16 empty] | 1.8812 | 102.4909
[16 empty] | 1.8813 | 102.4681
[16 empty] | 1.8814 | 102.4453
[16 empty] | 1.8815 | 102.4225
[16 empty] | 1.8816 | 102.3997
[16 empty] | 1.8817 | 102.3769
[16 empty] | 1.8818 | 102.3541
[16 empty] | 1.8818 | 102.3313
[16 empty] | 1.8819 | 102.3085
[16 empty] | 1.882 | 102.2857
[16 empty] | 1.8821 | 102.2629
[16 empty] | 1.8822 | 102.2402
[16 empty] | 1.8823 | 102.2174
[16 empty] | 1.8824 | 102.1946
[16 empty] | 1.8825 | 102.1718
[16 empty] | 1.8825 | 102.149
[16 empty] | 1.8826 | 102.1262
[16 empty] | 1.8827 | 102.1034
[16 empty] | 1.8828 | 102.0806
[16 empty] | 1.8829 | 102.0578
[16 empty] | 1.883 | 102.035
[16 empty] | 1.8831 | 102.0122
[16 empty] | 1.8832 | 101.9894
[16 empty] | 1.8832 | 101.9666
[16 empty] | 1.8833 | 101.9438
[16 empty] | 1.8834 | 101.921
[16 empty] | 1.8835 | 101.8982
[16 empty] | 1.8836 | 101.8755
[16 empty] | 1.8837 | 101.8527
[16 empty] | 1.8838 | 101.8299
[16 empty] | 1.8839 | 101.8071
[16 empty] | 1.8839 | 101.7843
[16 empty] | 1.884 | 101.7615
[16 empty] | 1.8841 | 101.7387
[16 empty] | 1.8842 | 101.7159
[16 empty] | 1.8843 | 101.6931
[16 empty] | 1.8844 | 101.6704
[16 empty] | 1.8845 | 101.6476
[16 empty] | 1.8846 | 101.6248
[16 empty] | 1.8846 | 101.602
[16 empty] | 1.8847 | 101.5792
[16 empty] | 1.8848 | 101.5564
[16 empty] | 1.8849 | 101.5336
[16 empty] | 1.885 | 101.5108
[16 empty] | 1.8851 | 101.4881
[16 empty] | 1.8852 | 101.4653
[16 empty] | 1.8853 | 101.4425
[16 empty] | 1.8853 | 101.4197
[16 empty] | 1.8854 | 101.3969
[16 empty] | 1.8855 | 101.3741
[16 empty] | 1.8856 | 101.3514
[16 empty] | 1.8857 | 101.3286
[16 empty] | 1.8858 | 101.3058
[16 empty] | 1.8859 | 101.283
[16 empty] | 1.886 | 101.2602
[16 empty] | 1.886 | 101.2374
[16 empty] | 1.8861 | 101.2147
[16 empty] | 1.8862 | 101.1919
[16 empty] | 1.8863 | 101.1691
[16 empty] | 1.8864 | 101.1463
[16 empty] | 1.8865 | 101.1235
[16 empty] | 1.8866 | 101.1008
[16 empty] | 1.8866 | 101.078
[16 empty] | 1.8867 | 101.0552
[16 empty] | 1.8868 | 101.0324
[16 empty] | 1.8869 | 101.0097
[16 empty] | 1.887 | 100.9869
[16 empty] | 1.8871 | 100.9641
[16 empty] | 1.8872 | 100.9413
[16 empty] | 1.8873 | 100.9186
[16 empty] | 1.8873 | 100.8958
[16 empty] | 1.8874 | 100.873
[16 empty] | 1.8875 | 100.8502
[16 empty] | 1.8876 | 100.8275
[16 empty] | 1.8877 | 100.8047
[16 empty] | 1.8878 | 100.7819
[16 empty] | 1.8879 | 100.7591
[16 empty] | 1.888 | 100.7364
[16 empty] | 1.888 | 100.7136
[16 empty] | 1.8881 | 100.6908
[16 empty] | 1.8882 | 100.6681
[16 empty] | 1.8883 | 100.6453
[16 empty] | 1.8884 | 100.6225
[16 empty] | 1.8885 | 100.5998
[16 empty] | 1.8886 | 100.577
[16 empty] | 1.8887 | 100.5542
[16 empty] | 1.8887 | 100.5315
[16 empty] | 1.8888 | 100.5087
[16 empty] | 1.8889 | 100.4859
[16 empty] | 1.889 | 100.4632
[16 empty] | 1.8891 | 100.4404
[16 empty] | 1.8892 | 100.4176
[16 empty] | 1.8893 | 100.3949
[16 empty] | 1.8894 | 100.3721
[16 empty] | 1.8894 | 100.3493
[16 empty] | 1.8895 | 100.3266
[16 empty] | 1.8896 | 100.3038
[16 empty] | 1.8897 | 100.2811
[16 empty] | 1.8898 | 100.2583
[16 empty] | 1.8899 | 100.2355
[16 empty] | 1.89 | 100.2128
[16 empty] | 1.8901 | 100.19
[16 empty] | 1.8901 | 100.1673
[16 empty] | 1.8902 | 100.1445
[16 empty] | 1.8903 | 100.1218
[16 empty] | 1.8904 | 100.099
[16 empty] | 1.8905 | 100.0763
[16 empty] | 1.8906 | 100.0535
[16 empty] | 1.8907 | 100.0307
[16 empty] | 1.8908 | 100.008
[16 empty] | 1.8908 | 99.9852
[16 empty] | 1.8909 | 99.9625
[16 empty] | 1.891 | 99.9397
[16 empty] | 1.8911 | 99.917
[16 empty] | 1.8912 | 99.8942
[16 empty] | 1.8913 | 99.8715
[16 empty] | 1.8914 | 99.8487
[16 empty] | 1.8915 | 99.826
[16 empty] | 1.8915 | 99.8033
[16 empty] | 1.8916 | 99.7805
[16 empty] | 1.8917 | 99.7578
[16 empty] | 1.8918 | 99.735
[16 empty] | 1.8919 | 99.7123
[16 empty] | 1.892 | 99.6895
[16 empty] | 1.8921 | 99.6668
[16 empty] | 1.8921 | 99.644
[16 empty] | 1.8922 | 99.6213
[16 empty] | 1.8923 | 99.5986
[16 empty] | 1.8924 | 99.5758
[16 empty] | 1.8925 | 99.5531
[16 empty] | 1.8926 | 99.5303
[16 empty] | 1.8927 | 99.5076
[16 empty] | 1.8928 | 99.4849
[16 empty] | 1.8928 | 99.4621
[16 empty] | 1.8929 | 99.4394
[16 empty] | 1.893 | 99.4167
[16 empty] | 1.8931 | 99.3939
[16 empty] | 1.8932 | 99.3712
[16 empty] | 1.8933 | 99.3485
[16 empty] | 1.8934 | 99.3257
[16 empty] | 1.8935 | 99.303
[16 empty] | 1.8935 | 99.2803
[16 empty] | 1.8936 | 99.2575
[16 empty] | 1.8937 | 99.2348
[16 empty] | 1.8938 | 99.2121
[16 empty] | 1.8939 | 99.1894
[16 empty] | 1.894 | 99.1666
[16 empty] | 1.8941 | 99.1439
[16 empty] | 1.8942 | 99.1212
[16 empty] | 1.8942 | 99.0985
[16 empty] | 1.8943 | 99.0757
[16 empty] | 1.8944 | 99.053
[16 empty] | 1.8945 | 99.0303
[16 empty] | 1.8946 | 99.0076
[16 empty] | 1.8947 | 98.9848
[16 empty] | 1.8948 | 98.9621
[16 empty] | 1.8949 | 98.9394
[16 empty] | 1.8949 | 98.9167
[16 empty] | 1.895 | 98.894
[16 empty] | 1.8951 | 98.8713
[16 empty] | 1.8952 | 98.8485
[16 empty] | 1.8953 | 98.8258
[16 empty] | 1.8954 | 98.8031
[16 empty] | 1.8955 | 98.7804
[16 empty] | 1.8956 | 98.7577
[16 empty] | 1.8956 | 98.735
[16 empty] | 1.8957 | 98.7123
[16 empty] | 1.8958 | 98.6896
[16 empty] | 1.8959 | 98.6668
[16 empty] | 1.896 | 98.6441
[16 empty] | 1.8961 | 98.6214
[16 empty] | 1.8962 | 98.5987
[16 empty] | 1.8963 | 98.576
[16 empty] | 1.8963 | 98.5533
[16 empty] | 1.8964 | 98.5306
[16 empty] | 1.8965 | 98.5079
[16 empty] | 1.8966 | 98.4852
[16 empty] | 1.8967 | 98.4625
[16 empty] | 1.8968 | 98.4398
[16 empty] | 1.8969 | 98.4171
[16 empty] | 1.897 | 98.3944
[16 empty] | 1.897 | 98.3717
[16 empty] | 1.8971 | 98.349
[16 empty] | 1.8972 | 98.3263
[16 empty] | 1.8973 | 98.3036
[16 empty] | 1.8974 | 98.2809
[16 empty] | 1.8975 | 98.2582
[16 empty] | 1.8976 | 98.2355
[16 empty] | 1.8977 | 98.2128
[16 empty] | 1.8977 | 98.1901
[16 empty] | 1.8978 | 98.1675
[16 empty] | 1.8979 | 98.1448
[16 empty] | 1.898 | 98.1221
[16 empty] | 1.8981 | 98.0994
[16 empty] | 1.8982 | 98.0767
[16 empty] | 1.8983 | 98.054
[16 empty] | 1.8983 | 98.0313
[16 empty] | 1.8984 | 98.0086
[16 empty] | 1.8985 | 97.986
[16 empty] | 1.8986 | 97.9633
[16 empty] | 1.8987 | 97.9406
[16 empty] | 1.8988 | 97.9179
[16 empty] | 1.8989 | 97.8952
[16 empty] | 1.899 | 97.8726
[16 empty] | 1.899 | 97.8499
[16 empty] | 1.8991 | 97.8272
[16 empty] | 1.8992 | 97.8045
[16 empty] | 1.8993 | 97.7819
[16 empty] | 1.8994 | 97.7592
[16 empty] | 1.8995 | 97.7365
[16 empty] | 1.8996 | 97.7138
[16 empty] | 1.8997 | 97.6912
[16 empty] | 1.8997 | 97.6685
[16 empty] | 1.8998 | 97.6458
[16 empty] | 1.8999 | 97.6232
[16 empty] | 1.9 | 97.6005
[16 empty] | 1.9001 | 97.5778
[16 empty] | 1.9002 | 97.5552
[16 empty] | 1.9003 | 97.5325
[16 empty] | 1.9004 | 97.5098
[16 empty] | 1.9004 | 97.4872
[16 empty] | 1.9005 | 97.4645
[16 empty] | 1.9006 | 97.4418
[16 empty] | 1.9007 | 97.4192
[16 empty] | 1.9008 | 97.3965
[16 empty] | 1.9009 | 97.3739
[16 empty] | 1.901 | 97.3512
[16 empty] | 1.9011 | 97.3286
[16 empty] | 1.9011 | 97.3059
[16 empty] | 1.9012 | 97.2832
[16 empty] | 1.9013 | 97.2606
[16 empty] | 1.9014 | 97.2379
[16 empty] | 1.9015 | 97.2153
[16 empty] | 1.9016 | 97.1926
[16 empty] | 1.9017 | 97.17
[16 empty] | 1.9018 | 97.1473
[16 empty] | 1.9018 | 97.1247
[16 empty] | 1.9019 | 97.102
[16 empty] | 1.902 | 97.0794
[16 empty] | 1.9021 | 97.0568
[16 empty] | 1.9022 | 97.0341
[16 empty] | 1.9023 | 97.0115
[16 empty] | 1.9024 | 96.9888
[16 empty] | 1.9025 | 96.9662
[16 empty] | 1.9025 | 96.9435
[16 empty] | 1.9026 | 96.9209
[16 empty] | 1.9027 | 96.8983
[16 empty] | 1.9028 | 96.8756
[16 empty] | 1.9029 | 96.853
[16 empty] | 1.903 | 96.8304
[16 empty] | 1.9031 | 96.8077
[16 empty] | 1.9032 | 96.7851
[16 empty] | 1.9032 | 96.7625
[16 empty] | 1.9033 | 96.7398
[16 empty] | 1.9034 | 96.7172
[16 empty] | 1.9035 | 96.6946
[16 empty] | 1.9036 | 96.672
[16 empty] | 1.9037 | 96.6493
[16 empty] | 1.9038 | 96.6267
[16 empty] | 1.9039 | 96.6041
[16 empty] | 1.9039 | 96.5815
[16 empty] | 1.904 | 96.5588
[16 empty] | 1.9041 | 96.5362
[16 empty] | 1.9042 | 96.5136
[16 empty] | 1.9043 | 96.491
[16 empty] | 1.9044 | 96.4684
[16 empty] | 1.9045 | 96.4457
[16 empty] | 1.9045 | 96.4231
[16 empty] | 1.9046 | 96.4005
[16 empty] | 1.9047 | 96.3779
[16 empty] | 1.9048 | 96.3553
[16 empty] | 1.9049 | 96.3327
[16 empty] | 1.905 | 96.3101
[16 empty] | 1.9051 | 96.2875
[16 empty] | 1.9052 | 96.2648
[16 empty] | 1.9052 | 96.2422
[16 empty] | 1.9053 | 96.2196
[16 empty] | 1.9054 | 96.197
[16 empty] | 1.9055 | 96.1744
[16 empty] | 1.9056 | 96.1518
[16 empty] | 1.9057 | 96.1292
[16 empty] | 1.9058 | 96.1066
[16 empty] | 1.9059 | 96.084
[16 empty] | 1.9059 | 96.0614
[16 empty] | 1.906 | 96.0388
[16 empty] | 1.9061 | 96.0162
[16 empty] | 1.9062 | 95.9936
[16 empty] | 1.9063 | 95.971
[16 empty] | 1.9064 | 95.9484
[16 empty] | 1.9065 | 95.9259
[16 empty] | 1.9066 | 95.9033
[16 empty] | 1.9066 | 95.8807
[16 empty] | 1.9067 | 95.8581
[16 empty] | 1.9068 | 95.8355
[16 empty] | 1.9069 | 95.8129
[16 empty] | 1.907 | 95.7903
[16 empty] | 1.9071 | 95.7677
[16 empty] | 1.9072 | 95.7452
[16 empty] | 1.9073 | 95.7226
[16 empty] | 1.9073 | 95.7
[16 empty] | 1.9074 | 95.6774
[16 empty] | 1.9075 | 95.6548
[16 empty] | 1.9076 | 95.6323
[16 empty] | 1.9077 | 95.6097
[16 empty] | 1.9078 | 95.5871
[16 empty] | 1.9079 | 95.5645
[16 empty] | 1.908 | 95.542
[16 empty] | 1.908 | 95.5194
[16 empty] | 1.9081 | 95.4968
[16 empty] | 1.9082 | 95.4742
[16 empty] | 1.9083 | 95.4517
[16 empty] | 1.9084 | 95.4291
[16 empty] | 1.9085 | 95.4065
[16 empty] | 1.9086 | 95.384
[16 empty] | 1.9087 | 95.3614
[16 empty] | 1.9087 | 95.3389
[16 empty] | 1.9088 | 95.3163
[16 empty] | 1.9089 | 95.2937
[16 empty] | 1.909 | 95.2712
[16 empty] | 1.9091 | 95.2486
[16 empty] | 1.9092 | 95.2261
[16 empty] | 1.9093 | 95.2035
[16 empty] | 1.9094 | 95.1809
[16 empty] | 1.9094 | 95.1584
[16 empty] | 1.9095 | 95.1358
[16 empty] | 1.9096 | 95.1133
[16 empty] | 1.9097 | 95.0907
[16 empty] | 1.9098 | 95.0682
[16 empty] | 1.9099 | 95.0456
[16 empty] | 1.91 | 95.0231
[16 empty] | 1.9101 | 95.0005
[16 empty] | 1.9101 | 94.978
[16 empty] | 1.9102 | 94.9555
[16 empty] | 1.9103 | 94.9329
[16 empty] | 1.9104 | 94.9104
[16 empty] | 1.9105 | 94.8878
[16 empty] | 1.9106 | 94.8653
[16 empty] | 1.9107 | 94.8428
[16 empty] | 1.9107 | 94.8202
[16 empty] | 1.9108 | 94.7977
[16 empty] | 1.9109 | 94.7752
[16 empty] | 1.911 | 94.7526
[16 empty] | 1.9111 | 94.7301
[16 empty] | 1.9112 | 94.7076
[16 empty] | 1.9113 | 94.685
[16 empty] | 1.9114 | 94.6625
[16 empty] | 1.9114 | 94.64
[16 empty] | 1.9115 | 94.6175
[16 empty] | 1.9116 | 94.5949
[16 empty] | 1.9117 | 94.5724
[16 empty] | 1.9118 | 94.5499
[16 empty] | 1.9119 | 94.5274
[16 empty] | 1.912 | 94.5048
[16 empty] | 1.9121 | 94.4823
[16 empty] | 1.9121 | 94.4598
[16 empty] | 1.9122 | 94.4373
[16 empty] | 1.9123 | 94.4148
[16 empty] | 1.9124 | 94.3923
[16 empty] | 1.9125 | 94.3698
[16 empty] | 1.9126 | 94.3472
[16 empty] | 1.9127 | 94.3247
[16 empty] | 1.9128 | 94.3022
[16 empty] | 1.9128 | 94.2797
[16 empty] | 1.9129 | 94.2572
[16 empty] | 1.913 | 94.2347
[16 empty] | 1.9131 | 94.2122
[16 empty] | 1.9132 | 94.1897
[16 empty] | 1.9133 | 94.1672
[16 empty] | 1.9134 | 94.1447
[16 empty] | 1.9135 | 94.1222
[16 empty] | 1.9135 | 94.0997
[16 empty] | 1.9136 | 94.0772
[16 empty] | 1.9137 | 94.0547
[16 empty] | 1.9138 | 94.0322
[16 empty] | 1.9139 | 94.0097
[16 empty] | 1.914 | 93.9872
[16 empty] | 1.9141 | 93.9647
[16 empty] | 1.9142 | 93.9423
[16 empty] | 1.9142 | 93.9198
[16 empty] | 1.9143 | 93.8973
[16 empty] | 1.9144 | 93.8748
[16 empty] | 1.9145 | 93.8523
[16 empty] | 1.9146 | 93.8298
[16 empty] | 1.9147 | 93.8073
[16 empty] | 1.9148 | 93.7849
[16 empty] | 1.9149 | 93.7624
[16 empty] | 1.9149 | 93.7399
[16 empty] | 1.915 | 93.7174
[16 empty] | 1.9151 | 93.695
[16 empty] | 1.9152 | 93.6725
[16 empty] | 1.9153 | 93.65
[16 empty] | 1.9154 | 93.6275
[16 empty] | 1.9155 | 93.6051
[16 empty] | 1.9156 | 93.5826
[16 empty] | 1.9156 | 93.5601
[16 empty] | 1.9157 | 93.5377
[16 empty] | 1.9158 | 93.5152
[16 empty] | 1.9159 | 93.4927
[16 empty] | 1.916 | 93.4703
[16 empty] | 1.9161 | 93.4478
[16 empty] | 1.9162 | 93.4253
[16 empty] | 1.9162 | 93.4029
[16 empty] | 1.9163 | 93.3804
[16 empty] | 1.9164 | 93.358
[16 empty] | 1.9165 | 93.3355
[16 empty] | 1.9166 | 93.3131
[16 empty] | 1.9167 | 93.2906
[16 empty] | 1.9168 | 93.2682
[16 empty] | 1.9169 | 93.2457
[16 empty] | 1.9169 | 93.2233
[16 empty] | 1.917 | 93.2008
[16 empty] | 1.9171 | 93.1784
[16 empty] | 1.9172 | 93.1559
[16 empty] | 1.9173 | 93.1335
[16 empty] | 1.9174 | 93.111
[16 empty] | 1.9175 | 93.0886
[16 empty] | 1.9176 | 93.0662
[16 empty] | 1.9176 | 93.0437
[16 empty] | 1.9177 | 93.0213
[16 empty] | 1.9178 | 92.9988
[16 empty] | 1.9179 | 92.9764
[16 empty] | 1.918 | 92.954
[16 empty] | 1.9181 | 92.9315
[16 empty] | 1.9182 | 92.9091
[16 empty] | 1.9183 | 92.8867
[16 empty] | 1.9183 | 92.8643
[16 empty] | 1.9184 | 92.8418
[16 empty] | 1.9185 | 92.8194
[16 empty] | 1.9186 | 92.797
[16 empty] | 1.9187 | 92.7745
[16 empty] | 1.9188 | 92.7521
[16 empty] | 1.9189 | 92.7297
[16 empty] | 1.919 | 92.7073
[16 empty] | 1.919 | 92.6849
[16 empty] | 1.9191 | 92.6625
[16 empty] | 1.9192 | 92.64
[16 empty] | 1.9193 | 92.6176
[16 empty] | 1.9194 | 92.5952
[16 empty] | 1.9195 | 92.5728
[16 empty] | 1.9196 | 92.5504
[16 empty] | 1.9197 | 92.528
[16 empty] | 1.9197 | 92.5056
[16 empty] | 1.9198 | 92.4832
[16 empty] | 1.9199 | 92.4608
[16 empty] | 1.92 | 92.4384
[16 empty] | 1.9201 | 92.416
[16 empty] | 1.9202 | 92.3936
[16 empty] | 1.9203 | 92.3712
[16 empty] | 1.9204 | 92.3488
[16 empty] | 1.9204 | 92.3264
[16 empty] | 1.9205 | 92.304
[16 empty] | 1.9206 | 92.2816
[16 empty] | 1.9207 | 92.2592
[16 empty] | 1.9208 | 92.2368
[16 empty] | 1.9209 | 92.2144
[16 empty] | 1.921 | 92.192
[16 empty] | 1.9211 | 92.1696
[16 empty] | 1.9211 | 92.1472
[16 empty] | 1.9212 | 92.1248
[16 empty] | 1.9213 | 92.1025
[16 empty] | 1.9214 | 92.0801
[16 empty] | 1.9215 | 92.0577
[16 empty] | 1.9216 | 92.0353
[16 empty] | 1.9217 | 92.0129
[16 empty] | 1.9218 | 91.9906
[16 empty] | 1.9218 | 91.9682
[16 empty] | 1.9219 | 91.9458
[16 empty] | 1.922 | 91.9234
[16 empty] | 1.9221 | 91.9011
[16 empty] | 1.9222 | 91.8787
[16 empty] | 1.9223 | 91.8563
[16 empty] | 1.9224 | 91.834
[16 empty] | 1.9224 | 91.8116
[16 empty] | 1.9225 | 91.7892
[16 empty] | 1.9226 | 91.7669
[16 empty] | 1.9227 | 91.7445
[16 empty] | 1.9228 | 91.7221
[16 empty] | 1.9229 | 91.6998
[16 empty] | 1.923 | 91.6774
[16 empty] | 1.9231 | 91.6551
[16 empty] | 1.9231 | 91.6327
[16 empty] | 1.9232 | 91.6103
[16 empty] | 1.9233 | 91.588
[16 empty] | 1.9234 | 91.5656
[16 empty] | 1.9235 | 91.5433
[16 empty] | 1.9236 | 91.5209
[16 empty] | 1.9237 | 91.4986
[16 empty] | 1.9238 | 91.4762
[16 empty] | 1.9238 | 91.4539
[16 empty] | 1.9239 | 91.4316
[16 empty] | 1.924 | 91.4092
[16 empty] | 1.9241 | 91.3869
[16 empty] | 1.9242 | 91.3645
[16 empty] | 1.9243 | 91.3422
[16 empty] | 1.9244 | 91.3199
[16 empty] | 1.9245 | 91.2975
[16 empty] | 1.9245 | 91.2752
[16 empty] | 1.9246 | 91.2529
[16 empty] | 1.9247 | 91.2305
[16 empty] | 1.9248 | 91.2082
[16 empty] | 1.9249 | 91.1859
[16 empty] | 1.925 | 91.1635
[16 empty] | 1.9251 | 91.1412
[16 empty] | 1.9252 | 91.1189
[16 empty] | 1.9252 | 91.0966
[16 empty] | 1.9253 | 91.0742
[16 empty] | 1.9254 | 91.0519
[16 empty] | 1.9255 | 91.0296
[16 empty] | 1.9256 | 91.0073
[16 empty] | 1.9257 | 90.985
[16 empty] | 1.9258 | 90.9627
[16 empty] | 1.9259 | 90.9403
[16 empty] | 1.9259 | 90.918
[16 empty] | 1.926 | 90.8957
[16 empty] | 1.9261 | 90.8734
[16 empty] | 1.9262 | 90.8511
[16 empty] | 1.9263 | 90.8288
[16 empty] | 1.9264 | 90.8065
[16 empty] | 1.9265 | 90.7842
[16 empty] | 1.9266 | 90.7619
[16 empty] | 1.9266 | 90.7396
[16 empty] | 1.9267 | 90.7173
[16 empty] | 1.9268 | 90.695
[16 empty] | 1.9269 | 90.6727
[16 empty] | 1.927 | 90.6504
[16 empty] | 1.9271 | 90.6281
[16 empty] | 1.9272 | 90.6058
[16 empty] | 1.9273 | 90.5835
[16 empty] | 1.9273 | 90.5612
[16 empty] | 1.9274 | 90.5389
[16 empty] | 1.9275 | 90.5166
[16 empty] | 1.9276 | 90.4944
[16 empty] | 1.9277 | 90.4721
[16 empty] | 1.9278 | 90.4498
[16 empty] | 1.9279 | 90.4275
[16 empty] | 1.928 | 90.4052
[16 empty] | 1.928 | 90.3829
[16 empty] | 1.9281 | 90.3607
[16 empty] | 1.9282 | 90.3384
[16 empty] | 1.9283 | 90.3161
[16 empty] | 1.9284 | 90.2938
[16 empty] | 1.9285 | 90.2716
[16 empty] | 1.9286 | 90.2493
[16 empty] | 1.9286 | 90.227
[16 empty] | 1.9287 | 90.2048
[16 empty] | 1.9288 | 90.1825
[16 empty] | 1.9289 | 90.1602
[16 empty] | 1.929 | 90.138
[16 empty] | 1.9291 | 90.1157
[16 empty] | 1.9292 | 90.0934
[16 empty] | 1.9293 | 90.0712
[16 empty] | 1.9293 | 90.0489
[16 empty] | 1.9294 | 90.0267
[16 empty] | 1.9295 | 90.0044
[16 empty] | 1.9296 | 89.9822
[16 empty] | 1.9297 | 89.9599
[16 empty] | 1.9298 | 89.9376
[16 empty] | 1.9299 | 89.9154
[16 empty] | 1.93 | 89.8931
[16 empty] | 1.93 | 89.8709
[16 empty] | 1.9301 | 89.8487
[16 empty] | 1.9302 | 89.8264
[16 empty] | 1.9303 | 89.8042
[16 empty] | 1.9304 | 89.7819
[16 empty] | 1.9305 | 89.7597
[16 empty] | 1.9306 | 89.7374
[16 empty] | 1.9307 | 89.7152
[16 empty] | 1.9307 | 89.693
[16 empty] | 1.9308 | 89.6707
[16 empty] | 1.9309 | 89.6485
[16 empty] | 1.931 | 89.6263
[16 empty] | 1.9311 | 89.604
[16 empty] | 1.9312 | 89.5818
[16 empty] | 1.9313 | 89.5596
[16 empty] | 1.9314 | 89.5374
[16 empty] | 1.9314 | 89.5151
[16 empty] | 1.9315 | 89.4929
[16 empty] | 1.9316 | 89.4707
[16 empty] | 1.9317 | 89.4485
[16 empty] | 1.9318 | 89.4263
[16 empty] | 1.9319 | 89.404
[16 empty] | 1.932 | 89.3818
[16 empty] | 1.9321 | 89.3596
[16 empty] | 1.9321 | 89.3374
[16 empty] | 1.9322 | 89.3152
[16 empty] | 1.9323 | 89.293
[16 empty] | 1.9324 | 89.2708
[16 empty] | 1.9325 | 89.2486
[16 empty] | 1.9326 | 89.2263
[16 empty] | 1.9327 | 89.2041
[16 empty] | 1.9328 | 89.1819
[16 empty] | 1.9328 | 89.1597
[16 empty] | 1.9329 | 89.1375
[16 empty] | 1.933 | 89.1153
[16 empty] | 1.9331 | 89.0931
[16 empty] | 1.9332 | 89.0709
[16 empty] | 1.9333 | 89.0487
[16 empty] | 1.9334 | 89.0266
[16 empty] | 1.9335 | 89.0044
[16 empty] | 1.9335 | 88.9822
[16 empty] | 1.9336 | 88.96
[16 empty] | 1.9337 | 88.9378
[16 empty] | 1.9338 | 88.9156
[16 empty] | 1.9339 | 88.8934
[16 empty] | 1.934 | 88.8712
[16 empty] | 1.9341 | 88.8491
[16 empty] | 1.9341 | 88.8269
[16 empty] | 1.9342 | 88.8047
[16 empty] | 1.9343 | 88.7825
[16 empty] | 1.9344 | 88.7603
[16 empty] | 1.9345 | 88.7382
[16 empty] | 1.9346 | 88.716
[16 empty] | 1.9347 | 88.6938
[16 empty] | 1.9348 | 88.6717
[16 empty] | 1.9348 | 88.6495
[16 empty] | 1.9349 | 88.6273
[16 empty] | 1.935 | 88.6051
[16 empty] | 1.9351 | 88.583
[16 empty] | 1.9352 | 88.5608
[16 empty] | 1.9353 | 88.5387
[16 empty] | 1.9354 | 88.5165
[16 empty] | 1.9355 | 88.4943
[16 empty] | 1.9355 | 88.4722
[16 empty] | 1.9356 | 88.45
[16 empty] | 1.9357 | 88.4279
[16 empty] | 1.9358 | 88.4057
[16 empty] | 1.9359 | 88.3836
[16 empty] | 1.936 | 88.3614
[16 empty] | 1.9361 | 88.3393
[16 empty] | 1.9362 | 88.3171
[16 empty] | 1.9362 | 88.295
[16 empty] | 1.9363 | 88.2728
[16 empty] | 1.9364 | 88.2507
[16 empty] | 1.9365 | 88.2285
[16 empty] | 1.9366 | 88.2064
[16 empty] | 1.9367 | 88.1843
[16 empty] | 1.9368 | 88.1621
[16 empty] | 1.9369 | 88.14
[16 empty] | 1.9369 | 88.1179
[16 empty] | 1.937 | 88.0957
[16 empty] | 1.9371 | 88.0736
[16 empty] | 1.9372 | 88.0515
[16 empty] | 1.9373 | 88.0293
[16 empty] | 1.9374 | 88.0072
[16 empty] | 1.9375 | 87.9851
[16 empty] | 1.9376 | 87.963
[16 empty] | 1.9376 | 87.9408
[16 empty] | 1.9377 | 87.9187
[16 empty] | 1.9378 | 87.8966
[16 empty] | 1.9379 | 87.8745
[16 empty] | 1.938 | 87.8524
[16 empty] | 1.9381 | 87.8302
[16 empty] | 1.9382 | 87.8081
[16 empty] | 1.9383 | 87.786
[16 empty] | 1.9383 | 87.7639
[16 empty] | 1.9384 | 87.7418
[16 empty] | 1.9385 | 87.7197
[16 empty] | 1.9386 | 87.6976
[16 empty] | 1.9387 | 87.6755
[16 empty] | 1.9388 | 87.6534
[16 empty] | 1.9389 | 87.6313
[16 empty] | 1.939 | 87.6092
[16 empty] | 1.939 | 87.5871
[16 empty] | 1.9391 | 87.565
[16 empty] | 1.9392 | 87.5429
[16 empty] | 1.9393 | 87.5208
[16 empty] | 1.9394 | 87.4987
[16 empty] | 1.9395 | 87.4766
[16 empty] | 1.9396 | 87.4545
[16 empty] | 1.9397 | 87.4324
[16 empty] | 1.9397 | 87.4103
[16 empty] | 1.9398 | 87.3882
[16 empty] | 1.9399 | 87.3661
[16 empty] | 1.94 | 87.3441
[16 empty] | 1.9401 | 87.322
[16 empty] | 1.9402 | 87.2999
[16 empty] | 1.9403 | 87.2778
[16 empty] | 1.9403 | 87.2557
[16 empty] | 1.9404 | 87.2337
[16 empty] | 1.9405 | 87.2116
[16 empty] | 1.9406 | 87.1895
[16 empty] | 1.9407 | 87.1674
[16 empty] | 1.9408 | 87.1454
[16 empty] | 1.9409 | 87.1233
[16 empty] | 1.941 | 87.1012
[16 empty] | 1.941 | 87.0792
[16 empty] | 1.9411 | 87.0571
[16 empty] | 1.9412 | 87.035
[16 empty] | 1.9413 | 87.013
[16 empty] | 1.9414 | 86.9909
[16 empty] | 1.9415 | 86.9689
[16 empty] | 1.9416 | 86.9468
[16 empty] | 1.9417 | 86.9248
[16 empty] | 1.9417 | 86.9027
[16 empty] | 1.9418 | 86.8806
[16 empty] | 1.9419 | 86.8586
[16 empty] | 1.942 | 86.8365
[16 empty] | 1.9421 | 86.8145
[16 empty] | 1.9422 | 86.7924
[16 empty] | 1.9423 | 86.7704
[16 empty] | 1.9424 | 86.7484
[16 empty] | 1.9424 | 86.7263
[16 empty] | 1.9425 | 86.7043
[16 empty] | 1.9426 | 86.6822
[16 empty] | 1.9427 | 86.6602
[16 empty] | 1.9428 | 86.6382
[16 empty] | 1.9429 | 86.6161
[16 empty] | 1.943 | 86.5941
[16 empty] | 1.9431 | 86.5721
[16 empty] | 1.9431 | 86.55
[16 empty] | 1.9432 | 86.528
[16 empty] | 1.9433 | 86.506
[16 empty] | 1.9434 | 86.4839
[16 empty] | 1.9435 | 86.4619
[16 empty] | 1.9436 | 86.4399
[16 empty] | 1.9437 | 86.4179
[16 empty] | 1.9438 | 86.3959
[16 empty] | 1.9438 | 86.3738
[16 empty] | 1.9439 | 86.3518
[16 empty] | 1.944 | 86.3298
[16 empty] | 1.9441 | 86.3078
[16 empty] | 1.9442 | 86.2858
[16 empty] | 1.9443 | 86.2638
[16 empty] | 1.9444 | 86.2417
[16 empty] | 1.9445 | 86.2197
[16 empty] | 1.9445 | 86.1977
[16 empty] | 1.9446 | 86.1757
[16 empty] | 1.9447 | 86.1537
[16 empty] | 1.9448 | 86.1317
[16 empty] | 1.9449 | 86.1097
[16 empty] | 1.945 | 86.0877
[16 empty] | 1.9451 | 86.0657
[16 empty] | 1.9452 | 86.0437
[16 empty] | 1.9452 | 86.0217
[16 empty] | 1.9453 | 85.9997
[16 empty] | 1.9454 | 85.9777
[16 empty] | 1.9455 | 85.9557
[16 empty] | 1.9456 | 85.9338
[16 empty] | 1.9457 | 85.9118
[16 empty] | 1.9458 | 85.8898
[16 empty] | 1.9459 | 85.8678
[16 empty] | 1.9459 | 85.8458
[16 empty] | 1.946 | 85.8238
[16 empty] | 1.9461 | 85.8019
[16 empty] | 1.9462 | 85.7799
[16 empty] | 1.9494 | 84.7438
[16 empty] | 1.9527 | 83.709
[16 empty] | 1.9559 | 82.6754
[16 empty] | 1.9592 | 81.643
[16 empty] | 1.9624 | 80.6119
[16 empty] | 1.9656 | 79.5821
[16 empty] | 1.9689 | 78.5535
[16 empty] | 1.9721 | 77.5261
[16 empty] | 1.9754 | 76.5001
[16 empty] | 1.9786 | 75.4753
[16 empty] | 1.9818 | 74.4518
[16 empty] | 1.9851 | 73.4296
[16 empty] | 1.9883 | 72.4086
[16 empty] | 1.9916 | 71.389
[16 empty] | 1.9948 | 70.3707
[16 empty] | 1.998 | 69.3536
[16 empty] | 2.0013 | 68.3379
[16 empty] | 2.0045 | 67.3235
[16 empty] | 2.0078 | 66.3103
[16 empty] | 2.011 | 65.2986
[16 empty] | 2.0142 | 64.2881
[16 empty] | 2.0175 | 63.279
[16 empty] | 2.0207 | 62.2712
[16 empty] | 2.024 | 61.2647
[16 empty] | 2.0272 | 60.2596
[16 empty] | 2.0304 | 59.2558
[16 empty] | 2.0337 | 58.2534
[16 empty] | 2.0369 | 57.2524
[16 empty] | 2.0402 | 56.2527
[16 empty] | 2.0434 | 55.2544
[16 empty] | 2.0467 | 54.2575
[16 empty] | 2.0499 | 53.2619
[16 empty] | 2.0531 | 52.2677
[16 empty] | 2.0564 | 51.275
[16 empty] | 2.0596 | 50.2836
[16 empty] | 2.0629 | 49.2936
[16 empty] | 2.0661 | 48.305
[16 empty] | 2.0693 | 47.3178
[16 empty] | 2.0726 | 46.3321
[16 empty] | 2.0758 | 45.3477
[16 empty] | 2.0791 | 44.3648
[16 empty] | 2.0823 | 43.3833
[16 empty] | 2.0855 | 42.4033
[16 empty] | 2.0888 | 41.4247
[16 empty] | 2.092 | 40.4475
[16 empty] | 2.0953 | 39.4718
[16 empty] | 2.0985 | 38.4975
[16 empty] | 2.1017 | 37.5247
[16 empty] | 2.105 | 36.5534
[16 empty] | 2.1082 | 35.5835
[16 empty] | 2.1115 | 34.6151
[16 empty] | 2.1147 | 33.6482
[16 empty] | 2.1179 | 32.6827
[16 empty] | 2.1212 | 31.7188
[16 empty] | 2.1244 | 30.7563
[16 empty] | 2.1277 | 29.7954
[16 empty] | 2.1309 | 28.8359
[16 empty] | 2.1341 | 27.878
[16 empty] | 2.1374 | 26.9216
[16 empty] | 2.1406 | 25.9666
[16 empty] | 2.1439 | 25.0133
[16 empty] | 2.1471 | 24.0614
[16 empty] | 2.1503 | 23.1111
[16 empty] | 2.1536 | 22.1623
[16 empty] | 2.1568 | 21.215
[16 empty] | 2.1601 | 20.2694
[16 empty] | 2.1633 | 19.3252
[16 empty] | 2.1665 | 18.3826
[16 empty] | 2.1698 | 17.4416
[16 empty] | 2.173 | 16.5022
[16 empty] | 2.1763 | 15.5643
[16 empty] | 2.1795 | 14.628
[16 empty] | 2.1827 | 13.6933
[16 empty] | 2.186 | 12.7601
[16 empty] | 2.1892 | 11.8286
[16 empty] | 2.1925 | 10.8987
[16 empty] | 2.1957 | 9.9703
[16 empty] | 2.1989 | 9.0436
[16 empty] | 2.2022 | 8.1185
[16 empty] | 2.2054 | 7.195
[16 empty] | 2.2087 | 6.2731
[16 empty] | 2.2119 | 5.3529
[16 empty] | 2.2152 | 4.4343
[16 empty] | 2.2184 | 3.5173
[16 empty] | 2.2216 | 2.6019
[16 empty] | 2.2249 | 1.6882
[16 empty] | 2.2281 | 0.7762
[16 empty] | 2.2314 | -0.13418
[16 empty] | 2.2346 | -1.0429
[16 empty] | 2.2378 | -1.95
[16 empty] | 2.2411 | -2.8553
[16 empty] | 2.2443 | -3.7591
[16 empty] | 2.2476 | -4.6611
[16 empty] | 2.2508 | -5.5614
[16 empty] | 2.254 | -6.4601
[16 empty] | 2.2573 | -7.357
[16 empty] | 2.2605 | -8.2523
[16 empty] | 2.2638 | -9.1459
[16 empty] | 2.267 | -10.0377
[16 empty] | 2.2702 | -10.9278